# ISO CLOSURE REGISTER | WOUND DRAINAGE

All primary applicable/conditional ISO 20697 clauses and related ISO normative annex methods • Evidence-based pending-status register

| PRODUCT Wound Drainage Set / flat perforated drain variants STANDARD ISO 20697:2018(E), corrected version 2018-09 | SOURCE WI WI-8-2-2-23 Rev.4, 25 Dec 2022 STATUS Controlled draft; no approval implied |
| --- | --- |

Evidence boundary: one official ISO PDF, one signed WI scan and two blank in-process XLSX forms were supplied. No executed records, master drawings, test reports, sterilization validation, risk file, or approved technical specifications were provided. A gap in supplied evidence is NOT evidence of a failed physical test.

# Assessment rules and closure categories

All 25 requirement groups are explicitly tracked, including feature-dependent requirements. No requirement is Verified Closed from blank source forms or proposed WI drafts. Documentary coverage is PROPOSED ONLY until controlled approvals. For product-conditional clauses, N/A is not assigned without an approved SKU-specific justification. Exact executed test IDs and quantitative company tolerances are pending.

# All-clause applicability, documentation and objective closure

| Requirement / clause / ISO page | ISO Says (technical paraphrase) | Related gap/change | Verification route | Required objective evidence | Actual verified evidence | Documentary coverage | Operational closure status | Outstanding action |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| REQ-01 4 PDF9 / ISO p3 | Safe access / functional drainage | GAP-002 | Design verification | Approved risk/technical applicability and corresponding signed validation, test, executed in-process or final evidence | Not provided / pending verification | Proposed WI sections and related source form(s); not approved | OPEN — evidence / technical input / approval | QA/Engineering: confirm Applicable; evidence not supplied |
| REQ-02 5.1 PDF9-10 / ISO pp3-4 | Risk management file | GAP-002 | Design/QA | Approved risk/technical applicability and corresponding signed validation, test, executed in-process or final evidence | Not provided / pending verification | Proposed WI sections and related source form(s); not approved | OPEN — evidence / technical input / approval | QA/Engineering: confirm Applicable; evidence not supplied |
| REQ-03 5.2 PDF10 / ISO p4 | Biocompatibility (ISO 10993-1) | GAP-003 | Design/QA | Approved risk/technical applicability and corresponding signed validation, test, executed in-process or final evidence | Not provided / pending verification | Proposed WI sections and related source form(s); not approved | OPEN — evidence / technical input / approval | QA/Engineering: confirm Applicable; evidence not supplied |
| REQ-04 5.3 PDF10 / ISO p4 | Detectability if risk requires | GAP-004 | Risk/qualification | Approved risk/technical applicability and corresponding signed validation, test, executed in-process or final evidence | Not provided / pending verification | Proposed WI sections and related source form(s); not approved | OPEN — applicability decision | QA/Engineering: confirm Conditional; decision pending |
| REQ-05 5.4 PDF10 / ISO p4 | Surface finish and condition for magnification | GAP-005 | In-process | Approved risk/technical applicability and corresponding signed validation, test, executed in-process or final evidence | Not provided / pending verification | Proposed WI sections and related source form(s); not approved | OPEN — evidence / technical input / approval | QA/Engineering: confirm Applicable; partial local controls |
| REQ-06 5.5.1-5.5.3 PDF10 / ISO p4 | OD designation; flat largest axis; effective length | GAP-006 | Dimensional IPC / design | Approved risk/technical applicability and corresponding signed validation, test, executed in-process or final evidence | Not provided / pending verification | Proposed WI sections and related source form(s); not approved | OPEN — evidence / technical input / approval | QA/Engineering: confirm Applicable; drawing missing |
| REQ-07 5.5.4 PDF11 / ISO p5 | Balloon nominal inflation volume | GAP-020 | Product configuration | Approved risk/technical applicability and corresponding signed validation, test, executed in-process or final evidence | Not provided / pending verification | Proposed WI sections and related source form(s); not approved | OPEN — applicability decision | QA/Engineering: confirm Conditional; balloon not verified |
| REQ-08 5.6 PDF11 / ISO p5 | Connection misconnection prevention | GAP-007 | Risk/qualification | Approved risk/technical applicability and corresponding signed validation, test, executed in-process or final evidence | Not provided / pending verification | Proposed WI sections and related source form(s); not approved | OPEN — evidence / technical input / approval | QA/Engineering: confirm Applicable to connectors; evidence pending |
| REQ-09 5.7 PDF12 / ISO p6 | MRI hazards if applicable | GAP-020 | Design / risk | Approved risk/technical applicability and corresponding signed validation, test, executed in-process or final evidence | Not provided / pending verification | Proposed WI sections and related source form(s); not approved | OPEN — applicability decision | QA/Engineering: confirm Conditional; intended use pending |
| REQ-10 5.8 PDF12 / ISO p6 | Sterilization SAL 10^-6 | GAP-008 | Validation / release | Approved risk/technical applicability and corresponding signed validation, test, executed in-process or final evidence | Not provided / pending verification | Proposed WI sections and related source form(s); not approved | OPEN — evidence / technical input / approval | QA/Engineering: confirm Applicable; evidence pending |
| REQ-11 6.1 PDF12 / ISO p6 | Kink performance risk-derived | GAP-009 | Qualification | Approved risk/technical applicability and corresponding signed validation, test, executed in-process or final evidence | Not provided / pending verification | Proposed WI sections and related source form(s); not approved | OPEN — evidence / technical input / approval | QA/Engineering: confirm Relevant to tube; criterion pending |
| REQ-12 6.2 + Annex B PDF12,19 / ISO pp6,13 | Metal corrosion risk / test | GAP-010 | Incoming/qualification | Approved risk/technical applicability and corresponding signed validation, test, executed in-process or final evidence | Not provided / pending verification | Proposed WI sections and related source form(s); not approved | OPEN — applicability decision | QA/Engineering: confirm Conditional for exposed trocar metal |
| REQ-13 6.3 + Annex C PDF12,20 / ISO pp6,14 | Vacuum collapse resistance | GAP-011 | Qualification | Approved risk/technical applicability and corresponding signed validation, test, executed in-process or final evidence | Not provided / pending verification | Proposed WI sections and related source form(s); not approved | OPEN — evidence / technical input / approval | QA/Engineering: confirm Suction configurations: applicable |
| REQ-14 6.4.1 + Table 1 + Annex D PDF12,21 / ISO pp6,15 | External connection tensile min 5/15 N by OD | GAP-012 | Qualification | Approved risk/technical applicability and corresponding signed validation, test, executed in-process or final evidence | Not provided / pending verification | Proposed WI sections and related source form(s); not approved | OPEN — evidence / technical input / approval | QA/Engineering: confirm Applicable external connections |
| REQ-15 6.4.2 + Table 2 + Annex E PDF13,22-23 / ISO pp7,16-17 | Tube/junction tensile min 10/20 N by OD | GAP-013 | Qualification | Approved risk/technical applicability and corresponding signed validation, test, executed in-process or final evidence | Not provided / pending verification | Proposed WI sections and related source form(s); not approved | OPEN — evidence / technical input / approval | QA/Engineering: confirm Applicable tube/junctions |
| REQ-16 6.5 + Annex F PDF13,24-25 / ISO pp7,18-19 | Collection impact no leak; vacuum loss <=2% | GAP-014 | Qualification | Approved risk/technical applicability and corresponding signed validation, test, executed in-process or final evidence | Not provided / pending verification | Proposed WI sections and related source form(s); not approved | OPEN — evidence / technical input / approval | QA/Engineering: confirm Collection/bellows: applicable |
| REQ-17 6.6 + Annex G PDF13,26-27 / ISO pp7,20-21 | Claimed flow threshold >=80% or >=90% | GAP-015 | Conditional qualification | Approved risk/technical applicability and corresponding signed validation, test, executed in-process or final evidence | Not provided / pending verification | Proposed WI sections and related source form(s); not approved | OPEN — evidence / technical input / approval | QA/Engineering: confirm Flow claim not provided |
| REQ-18 6.7 + Annex H PDF13,28-29 / ISO pp7,22-23 | Non-balloon retention performance | GAP-020 | Risk/qualification | Approved risk/technical applicability and corresponding signed validation, test, executed in-process or final evidence | Not provided / pending verification | Proposed WI sections and related source form(s); not approved | OPEN — applicability decision | QA/Engineering: confirm Only if retention mechanism present |
| REQ-19 6.8 + Annex I PDF13,30-32 / ISO pp7,24-26 | Balloon safety | GAP-020 | Qualification | Approved risk/technical applicability and corresponding signed validation, test, executed in-process or final evidence | Not provided / pending verification | Proposed WI sections and related source form(s); not approved | OPEN — applicability decision | QA/Engineering: confirm Only if balloon present |
| REQ-20 6.9 + Annex J/K PDF14,33-36 / ISO pp8,27-30 | Balloon deflation recovery / profile | GAP-020 | Qualification | Approved risk/technical applicability and corresponding signed validation, test, executed in-process or final evidence | Not provided / pending verification | Proposed WI sections and related source form(s); not approved | OPEN — applicability decision | QA/Engineering: confirm Only if appropriate balloon present |
| REQ-21 6.10 + Annex L PDF14,37-40 / ISO pp8,31-34 | Inflated balloon traction | GAP-020 | Qualification | Approved risk/technical applicability and corresponding signed validation, test, executed in-process or final evidence | Not provided / pending verification | Proposed WI sections and related source form(s); not approved | OPEN — evidence / technical input / approval | QA/Engineering: confirm Only for applicable balloon geometry |
| REQ-22 6.11 + Annex M PDF14,41-42 / ISO pp8,35-36 | No air leak at max vacuum | GAP-016 | Qualification / IPC | Approved risk/technical applicability and corresponding signed validation, test, executed in-process or final evidence | Not provided / pending verification | Proposed WI sections and related source form(s); not approved | OPEN — evidence / technical input / approval | QA/Engineering: confirm Applicable suction assembly |
| REQ-23 7.1 PDF15 / ISO p9 | Mark durability/legibility | GAP-018 | Packaging / qualification | Approved risk/technical applicability and corresponding signed validation, test, executed in-process or final evidence | Not provided / pending verification | Proposed WI sections and related source form(s); not approved | OPEN — evidence / technical input / approval | QA/Engineering: confirm Applicable |
| REQ-24 7.2 PDF15 / ISO p9 | Packaging marking required fields | GAP-017; GAP-018 | Packaging review | Approved risk/technical applicability and corresponding signed validation, test, executed in-process or final evidence | Not provided / pending verification | Proposed WI sections and related source form(s); not approved | OPEN — evidence / technical input / approval | QA/Engineering: confirm Applicable per feature |
| REQ-25 7.3 PDF15-16 / ISO pp9-10 | IFU provided: minimum contents | GAP-018 | Document control | Approved risk/technical applicability and corresponding signed validation, test, executed in-process or final evidence | Not provided / pending verification | Proposed WI sections and related source form(s); not approved | OPEN — applicability decision | QA/Engineering: confirm Conditional upon separate IFU |

# Gap-by-gap closure status and manager authorization

| Gap | Priority | Evidence to obtain | Technical input / approver | Status |
| --- | --- | --- | --- | --- |
| GAP-001 | P2 | Approved document index, WI/form cross-reference, master list entry, documented revision decisions. No approved master list supplied. | QC Manager / QA / Engineering, per approved responsibility matrix | OPEN — evidence / inputs / approvals |
| GAP-002 | P1 | Approved risk-management file and report, design verification matrix, product configurations and design claim evidence; none provided. | QC Manager / QA / Engineering, per approved responsibility matrix | OPEN — evidence / inputs / approvals |
| GAP-003 | P1 | Approved biological evaluation report/plan, material specs, supplier traceability and change assessments; not supplied. | QC Manager / QA / Engineering, per approved responsibility matrix | OPEN — evidence / inputs / approvals |
| GAP-004 | P1 | Detectability risk decision, approved drawing and performance qualification evidence, not supplied. | QC Manager / QA / Engineering, per approved responsibility matrix | OPEN — evidence / inputs / approvals |
| GAP-005 | P2 | Approved visual defect criteria, risk-based magnification decision, signed sample records/defect disposition; not provided. | QC Manager / QA / Engineering, per approved responsibility matrix | OPEN — evidence / inputs / approvals |
| GAP-006 | P2 | Approved size-specific drawings, calibrated measuring equipment and real measurements. OPEN TECHNICAL INPUT REQUIRED for missing drawing-specific limits and configurations. | QC Manager / QA / Engineering, per approved responsibility matrix | OPEN — evidence / inputs / approvals |
| GAP-007 | P1 | Approved interface specification and ISO 80369-1-based risk assessment; not supplied. | QC Manager / QA / Engineering, per approved responsibility matrix | OPEN — evidence / inputs / approvals |
| GAP-008 | P1 | Validated sterilization file, approved SAL rationale and batch cycle/release records; WI-7-5-1-1 not supplied. Operational closure unavailable. | QC Manager / QA / Engineering, per approved responsibility matrix | OPEN — evidence / inputs / approvals |
| GAP-009 | P2 | Approved clinical/risk justification, kink method/criteria and executed evidence; OPEN TECHNICAL INPUT REQUIRED. | QC Manager / QA / Engineering, per approved responsibility matrix | OPEN — evidence / inputs / approvals |
| GAP-010 | P2 | Exposed-metal risk assessment, Annex B test report and traceable incoming verification; no executed reports provided. | QC Manager / QA / Engineering, per approved responsibility matrix | OPEN — evidence / inputs / approvals |
| GAP-011 | P1 | Approved max negative pressure, executed Annex C report, equipment calibration and risk-approved verification frequency. OPEN TECHNICAL INPUT REQUIRED for pressure. | QC Manager / QA / Engineering, per approved responsibility matrix | OPEN — evidence / inputs / approvals |
| GAP-012 | P1 | Annex D signed tensile results, connector/size matrix, validated bond process and approved monitoring plan; none provided. | QC Manager / QA / Engineering, per approved responsibility matrix | OPEN — evidence / inputs / approvals |
| GAP-013 | P1 | Annex E report, bonding process qualification and change-control scope; not supplied. | QC Manager / QA / Engineering, per approved responsibility matrix | OPEN — evidence / inputs / approvals |
| GAP-014 | P1 | Impact qualification test reports and controlled design/lot change coverage. Not supplied. | QC Manager / QA / Engineering, per approved responsibility matrix | OPEN — evidence / inputs / approvals |
| GAP-015 | P2 | Manufacturer claims/IFU, Annex G test evidence or approved N/A assessment; OPEN TECHNICAL INPUT REQUIRED on flow claim. | QC Manager / QA / Engineering, per approved responsibility matrix | OPEN — evidence / inputs / approvals |
| GAP-016 | P1 | Approved maximum pressure, Annex M executed qualification, validated in-process manual test parameters and records; OPEN TECHNICAL INPUT REQUIRED for routine pressure/time. | QC Manager / QA / Engineering, per approved responsibility matrix | OPEN — evidence / inputs / approvals |
| GAP-017 | P2 | Approved bellows spec, volume tolerance, stable-vacuum qualification and approved artwork; no data provided. | QC Manager / QA / Engineering, per approved responsibility matrix | OPEN — evidence / inputs / approvals |
| GAP-018 | P2 | Approved artworks/IFUs, packaging controls/validation and completed lot packaging records F-823-i4-2; not supplied. | QC Manager / QA / Engineering, per approved responsibility matrix | OPEN — evidence / inputs / approvals |
| GAP-019 | P2 | Approved TF-23-PS/TF-23-PDBM, drawings, sampling/control plan, instrument calibration and executed signed in-process records; none supplied. OPEN TECHNICAL INPUT REQUIRED for unspecified tolerances and tests. | QC Manager / QA / Engineering, per approved responsibility matrix | OPEN — evidence / inputs / approvals |
| GAP-020 | P2 | Approved BOM/drawing, intended use and risk assessment with signed applicability verdict; not supplied. | QC Manager / QA / Engineering, per approved responsibility matrix | OPEN — evidence / inputs / approvals |

# Evidence-based closure sign-off (blank by design)

| Review gate | Decision / identified actual source | Name / position | Date / signature |
| --- | --- | --- | --- |
| Approved product applicability and clause disposition | ________________________ | ________________ | ________________ |
| Controlled WI + two in-process form revisions effective | ________________________ | ________________ | ________________ |
| Missing limits / max pressure / claims approved | ________________________ | ________________ | ________________ |
| Qualification and validation evidence independently verified | ________________________ | ________________ | ________________ |
| QC Manager release of audit findings only (not product release) | ________________________ | ________________ | ________________ |

# Source traceability / supplied document register

| ID | Source | Control status / observation |
| --- | --- | --- |
| SRC-ISO | ISO 20697:2018(E), corrected version 2018-09 (44-page PDF, numbered ISO pages 1-37; ISO page = PDF page - 6 from PDF7 onward) | Licensed source; technical paraphrases rather than long copyrighted extracts; no signed applicability matrix |
| SRC-WI | WI-8-2-2-23 Rev.4, 25 Dec 2022; supplied 7-page signed scanned PDF | Controlled Copy stamp visible; master-list/effective status not independently verified |
| SRC-IP1 | F-823-i12 Rev.4, 23 Nov 2025 (one-sheet XLSX; rows 11-27) | Blank in-process production form; reference discrepancy at F4; source format intact |
| SRC-IP2 | F-823-i211 Rev.4, 06 Dec 2025 (one-sheet XLSX; rows 10-24) | Blank flat-drain assembly form; internal sheet tab says Rev.0 while Q3 shows 4 |

Citation convention: ISO PDF page numbers count the cover as PDF p.1; published numbered text begins at PDF p.7 = printed p.1. WI pages are numbered 1-7 on the document. XLSX row numbers are native sheet row indices. All pending evidence references are explicitly unverified.