# WOUND DRAINAGE | MASTER GAP ANALYSIS

ISO requirement-to-WI-to-In-Process control reconciliation • QC Manager briefing

| PRODUCT Wound Drainage Set / flat perforated drain variants STANDARD ISO 20697:2018(E), corrected version 2018-09 | SOURCE WI WI-8-2-2-23 Rev.4, 25 Dec 2022 STATUS Controlled draft; no approval implied |
| --- | --- |

Evidence boundary: one official ISO PDF, one signed WI scan and two blank in-process XLSX forms were supplied. No executed records, master drawings, test reports, sterilization validation, risk file, or approved technical specifications were provided. A gap in supplied evidence is NOT evidence of a failed physical test.

# 1. Executive summary

Product: Wound Drainage Set and flat perforated silicone drain variants. Standard: ISO 20697:2018(E), corrected version 2018-09. Review covered 20 distinct documentary/control/evidence gaps: P1=10; P2=10; P3=0. Gaps record omissions or missing accessible objective evidence, NOT confirmed product failures.

Management decision: revision of WI and both in-process records is recommended, together with controlled links to separate design qualification, pressure tests, tensile verification, sterilization validation, supplier evidence and packaging/release records. These two forms are NOT final inspection reports. A revised draft must not be issued as effective until controlled change, risk review and approvals are complete.

Highest-risk unresolved evidence: risk file/biocompatibility, maximum negative pressure (Annex C/M), quantitative tensile force (Annex D/E), impact (Annex F), validated sterilization at SAL 10^-6, and connection/suction performance. Product-specific dimensional tolerances and optional-feature applicability remain OPEN TECHNICAL INPUT REQUIRED.

# 2. Sources and assessment limits

# Source traceability / supplied document register

| ID | Source | Control status / observation |
| --- | --- | --- |
| SRC-ISO | ISO 20697:2018(E), corrected version 2018-09 (44-page PDF, numbered ISO pages 1-37; ISO page = PDF page - 6 from PDF7 onward) | Licensed source; technical paraphrases rather than long copyrighted extracts; no signed applicability matrix |
| SRC-WI | WI-8-2-2-23 Rev.4, 25 Dec 2022; supplied 7-page signed scanned PDF | Controlled Copy stamp visible; master-list/effective status not independently verified |
| SRC-IP1 | F-823-i12 Rev.4, 23 Nov 2025 (one-sheet XLSX; rows 11-27) | Blank in-process production form; reference discrepancy at F4; source format intact |
| SRC-IP2 | F-823-i211 Rev.4, 06 Dec 2025 (one-sheet XLSX; rows 10-24) | Blank flat-drain assembly form; internal sheet tab says Rev.0 while Q3 shows 4 |

Citation convention: ISO PDF page numbers count the cover as PDF p.1; published numbered text begins at PDF p.7 = printed p.1. WI pages are numbered 1-7 on the document. XLSX row numbers are native sheet row indices. All pending evidence references are explicitly unverified.

# 3. Applicability and verified ISO requirement inventory

| Requirement | ISO clause | ISO-derived control summary | ISO source | Linked gap | Verification route / status |
| --- | --- | --- | --- | --- | --- |
| REQ-01 | 4 | Safe access / functional drainage | PDF9 / ISO p3 | GAP-002 | Design verification; Applicable; evidence not supplied |
| REQ-02 | 5.1 | Risk management file | PDF9-10 / ISO pp3-4 | GAP-002 | Design/QA; Applicable; evidence not supplied |
| REQ-03 | 5.2 | Biocompatibility (ISO 10993-1) | PDF10 / ISO p4 | GAP-003 | Design/QA; Applicable; evidence not supplied |
| REQ-04 | 5.3 | Detectability if risk requires | PDF10 / ISO p4 | GAP-004 | Risk/qualification; Conditional; decision pending |
| REQ-05 | 5.4 | Surface finish and condition for magnification | PDF10 / ISO p4 | GAP-005 | In-process; Applicable; partial local controls |
| REQ-06 | 5.5.1-5.5.3 | OD designation; flat largest axis; effective length | PDF10 / ISO p4 | GAP-006 | Dimensional IPC / design; Applicable; drawing missing |
| REQ-07 | 5.5.4 | Balloon nominal inflation volume | PDF11 / ISO p5 | GAP-020 | Product configuration; Conditional; balloon not verified |
| REQ-08 | 5.6 | Connection misconnection prevention | PDF11 / ISO p5 | GAP-007 | Risk/qualification; Applicable to connectors; evidence pending |
| REQ-09 | 5.7 | MRI hazards if applicable | PDF12 / ISO p6 | GAP-020 | Design / risk; Conditional; intended use pending |
| REQ-10 | 5.8 | Sterilization SAL 10^-6 | PDF12 / ISO p6 | GAP-008 | Validation / release; Applicable; evidence pending |
| REQ-11 | 6.1 | Kink performance risk-derived | PDF12 / ISO p6 | GAP-009 | Qualification; Relevant to tube; criterion pending |
| REQ-12 | 6.2 + Annex B | Metal corrosion risk / test | PDF12,19 / ISO pp6,13 | GAP-010 | Incoming/qualification; Conditional for exposed trocar metal |
| REQ-13 | 6.3 + Annex C | Vacuum collapse resistance | PDF12,20 / ISO pp6,14 | GAP-011 | Qualification; Suction configurations: applicable |
| REQ-14 | 6.4.1 + Table 1 + Annex D | External connection tensile min 5/15 N by OD | PDF12,21 / ISO pp6,15 | GAP-012 | Qualification; Applicable external connections |
| REQ-15 | 6.4.2 + Table 2 + Annex E | Tube/junction tensile min 10/20 N by OD | PDF13,22-23 / ISO pp7,16-17 | GAP-013 | Qualification; Applicable tube/junctions |
| REQ-16 | 6.5 + Annex F | Collection impact no leak; vacuum loss <=2% | PDF13,24-25 / ISO pp7,18-19 | GAP-014 | Qualification; Collection/bellows: applicable |
| REQ-17 | 6.6 + Annex G | Claimed flow threshold >=80% or >=90% | PDF13,26-27 / ISO pp7,20-21 | GAP-015 | Conditional qualification; Flow claim not provided |
| REQ-18 | 6.7 + Annex H | Non-balloon retention performance | PDF13,28-29 / ISO pp7,22-23 | GAP-020 | Risk/qualification; Only if retention mechanism present |
| REQ-19 | 6.8 + Annex I | Balloon safety | PDF13,30-32 / ISO pp7,24-26 | GAP-020 | Qualification; Only if balloon present |
| REQ-20 | 6.9 + Annex J/K | Balloon deflation recovery / profile | PDF14,33-36 / ISO pp8,27-30 | GAP-020 | Qualification; Only if appropriate balloon present |
| REQ-21 | 6.10 + Annex L | Inflated balloon traction | PDF14,37-40 / ISO pp8,31-34 | GAP-020 | Qualification; Only for applicable balloon geometry |
| REQ-22 | 6.11 + Annex M | No air leak at max vacuum | PDF14,41-42 / ISO pp8,35-36 | GAP-016 | Qualification / IPC; Applicable suction assembly |
| REQ-23 | 7.1 | Mark durability/legibility | PDF15 / ISO p9 | GAP-018 | Packaging / qualification; Applicable |
| REQ-24 | 7.2 | Packaging marking required fields | PDF15 / ISO p9 | GAP-017; GAP-018 | Packaging review; Applicable per feature |
| REQ-25 | 7.3 | IFU provided: minimum contents | PDF15-16 / ISO pp9-10 | GAP-018 | Document control; Conditional upon separate IFU |

# 4. Initial findings | priority

| Gap | Priority | Finding | Rationale for priority |
| --- | --- | --- | --- |
| GAP-001 | P2 | Controlled document reference conflict | Meaningful verification, traceability or documentation/control weakness |
| GAP-002 | P1 | Intended performance and risk-management traceability | Patient safety / critical functional performance or critical conformity evidence |
| GAP-003 | P1 | Biological evaluation evidence not established | Patient safety / critical functional performance or critical conformity evidence |
| GAP-004 | P1 | Detectability / X-ray claim verification unclear | Patient safety / critical functional performance or critical conformity evidence |
| GAP-005 | P2 | Surface acceptance insufficiently defined | Meaningful verification, traceability or documentation/control weakness |
| GAP-006 | P2 | Size, diameter and non-round profile traceability | Meaningful verification, traceability or documentation/control weakness |
| GAP-007 | P1 | Connector misconnection safety unverified | Patient safety / critical functional performance or critical conformity evidence |
| GAP-008 | P1 | Sterilization qualification and release linkage not verified | Patient safety / critical functional performance or critical conformity evidence |
| GAP-009 | P2 | Kink criterion and evidence not defined by risk | Meaningful verification, traceability or documentation/control weakness |
| GAP-010 | P2 | Trocar corrosion WI omits specified normative procedure | Meaningful verification, traceability or documentation/control weakness |
| GAP-011 | P1 | Vacuum deformation lacks defined pressure and method | Patient safety / critical functional performance or critical conformity evidence |
| GAP-012 | P1 | External connector tensile force not quantitatively shown | Patient safety / critical functional performance or critical conformity evidence |
| GAP-013 | P1 | Catheter tubular/junction tensile criteria not determined | Patient safety / critical functional performance or critical conformity evidence |
| GAP-014 | P1 | Collection device impact/vacuum retention verification missing | Patient safety / critical functional performance or critical conformity evidence |
| GAP-015 | P2 | Conditional claimed flow performance not traceable | Meaningful verification, traceability or documentation/control weakness |
| GAP-016 | P1 | Manual bellows leak check does not establish Annex M vacuum leakage conformity | Patient safety / critical functional performance or critical conformity evidence |
| GAP-017 | P2 | Capacity, pre-evacuated vacuum claim and label criteria unlinked | Meaningful verification, traceability or documentation/control weakness |
| GAP-018 | P2 | Packaging label/IFU content review evidence incomplete | Meaningful verification, traceability or documentation/control weakness |
| GAP-019 | P2 | In-process characteristics lack objective result capture / approved technical limits | Meaningful verification, traceability or documentation/control weakness |
| GAP-020 | P2 | MRI and optional feature applicability not formally recorded | Meaningful verification, traceability or documentation/control weakness |

# 5. MASTER GAP & REVISION TABLE

Mandatory seven-column headings are retained exactly as requested. IMPORTANT: columns labelled "Existing Inspection Text" and "Revised Inspection Text" refer to the two supplied IN-PROCESS CONTROL FORMS ONLY. No finished-product Inspection & Test Report was supplied. Proposed separate qualification evidence is linked, not misrepresented as hourly in-process testing.

| Gap / Our Problem | ISO Says | Existing WI Text | Revised WI Text | Existing Inspection Text | Revised Inspection Text | Evidence (ISO Closure) |
| --- | --- | --- | --- | --- | --- | --- |
| GAP-001 ¦ P2 Controlled document reference conflict Problem: existing controls do not establish fully objective requirement and/or executed traceable proof for the specific characteristic. | Document-control reconciliation; ISO 20697 does not assign internal form or WI numbers. | WI p.7 §14 lists F-823-i12. F-823-i12 Rev.4 cell F4 cites WI-8-2-2-23-1 / TF-23-PS / TF-23-PDBM; the supplied WI header is WI-8-2-2-23 Rev.4. F-823-i211 Rev.4 cell D3 cites WI-8-2-2-23; its worksheet tab is labelled Rev.0 although Q3 states Rev.4. | §17 Document Control (proposed): Before issue or use, Document Control shall reconcile WI, drawing, process specification and in-process record identifiers and revisions against the approved master list. Any nonmatching reference or worksheet-tab designation shall be logged, technically dispositioned and corrected through controlled revision. Do not assume -23 and -23-1 are interchangeable. | F-823-i12: F4 WI suffix -23-1; F-823-i211: D3 WI -23, Q3 Rev.4, tab name Rev.0. Both are in-process records, not final inspection reports. | Form headers: Controlled WI ID/rev; applicable TF-23-PS and TF-23-PDBM rev; drawing ID/rev; form ID/rev; effective date; controlled master-list confirmation; discrepancy disposition reference. Preserve actual form numbers until approved change. | Required Evidence: Approved document index, WI/form cross-reference, master list entry, documented revision decisions. No approved master list supplied. Verification / Control Route: Document Control Evidence Reference: Pending / not supplied Closure Criteria: approved applicable control, source-linked execution evidence and authorized disposition. Current Closure Status: OPEN — evidence / technical input / approval; draft wording is not operational closure. |
| GAP-002 ¦ P1 Intended performance and risk-management traceability Problem: existing controls do not establish fully objective requirement and/or executed traceable proof for the specific characteristic. | ISO §4 (PDF9 / printed3) requires safe intended access, maintained drainage and relevant retention performance; §5.1 (PDF9-10 / printed3-4) requires risk management and risk analysis, verified using a risk-management file. | WI p.1 §§1-3 describes wound drainage and quality objective. WI pp.2-6 lists checks but gives no risk-file link, validated performance claims or product-configuration applicability matrix. | §§2,5.1 (proposed): For each wound drainage configuration, the approved device master specification shall define intended purpose, components, suction mode, relevant performance claims and permitted variants. QA/Engineering shall maintain the approved risk-management file and trace identified hazards to verification and process controls. QC shall use only released drawings/control plans matched to the work order. | F-823-i12 rows 11-27 and F-823-i211 rows 10-24 contain local production checkpoints, not a documented intended-performance/risk-verification record. | Header / linkage only: product code; configuration; intended-use/specification revision; risk-file ID and controlled verification evidence reference. Keep risk evaluation and design validation outside routine in-process records. | Required Evidence: Approved risk-management file and report, design verification matrix, product configurations and design claim evidence; none provided. Verification / Control Route: Design/QA qualification Evidence Reference: Pending / not supplied Closure Criteria: approved applicable control, source-linked execution evidence and authorized disposition. Current Closure Status: OPEN — evidence / technical input / approval; draft wording is not operational closure. |
| GAP-003 ¦ P1 Biological evaluation evidence not established Problem: existing controls do not establish fully objective requirement and/or executed traceable proof for the specific characteristic. | ISO §5.2 (PDF10 / printed4) requires freedom from biological hazard using appropriate evaluation/testing according to ISO 10993-1. | WI pp.1-7 has no specified biological evaluation documentation pathway. The listed incoming checks at p.1 §5.1.1 are visual/commercial, not biological evaluation. | §5.2 (proposed): QA shall verify an approved biological evaluation covering the finished device and relevant patient-contact materials, processing and sterilization in accordance with the applicable ISO 10993-1 evaluation process. In-process QC shall verify material/component identity and change status against released specifications; any unapproved substitution shall be held. | F-823-i12 row 24 states "Correct specification / As per requirement"; F-823-i211 rows 12 and 19 identify silicone components, without biocompatibility evidence IDs. | Material identification / approval: drawing rev, component lots, authorized material IDs, approved change status, evaluation-file reference. Do not add biological laboratory tests to every hourly checklist. | Required Evidence: Approved biological evaluation report/plan, material specs, supplier traceability and change assessments; not supplied. Verification / Control Route: Design/QA qualification Evidence Reference: Pending / not supplied Closure Criteria: approved applicable control, source-linked execution evidence and authorized disposition. Current Closure Status: OPEN — evidence / technical input / approval; draft wording is not operational closure. |
| GAP-004 ¦ P1 Detectability / X-ray claim verification unclear Problem: existing controls do not establish fully objective requirement and/or executed traceable proof for the specific characteristic. | ISO §5.3 (PDF10 / printed4): if required by risk assessment, the catheter/effective length must be detectable by X-ray or another defined means. | WI p.3 §6.2 specifies a dark-blue visible X-ray line on the Redon tube; p.3 §6.3 and p.6 §10.5 mention X-ray line/ASTM F640. The requirement is not matched to an approved detectability risk decision and qualified evidence. | §10.5 (proposed): For each SKU requiring imaging detectability under the approved risk assessment, verify the specified X-ray marker is continuous and correctly positioned during production. QA shall retain performance verification under the applicable controlled method and device configuration. If detectability is not required for a configuration, document justified N/A. | F-823-i211 row 19: "Visible X-ray line" for round silicone tube. F-823-i12 has no dedicated detectability checkpoint. | Flat-drain assembly checkpoint: X-ray marker continuity / position as approved drawing; actual observation Pass/Fail and evidence ID. Qualification linkage: imaging detectability test report and risk disposition. | Required Evidence: Detectability risk decision, approved drawing and performance qualification evidence, not supplied. Verification / Control Route: In-process plus qualification Evidence Reference: Pending / not supplied Closure Criteria: approved applicable control, source-linked execution evidence and authorized disposition. Current Closure Status: OPEN — evidence / technical input / approval; draft wording is not operational closure. |
| GAP-005 ¦ P2 Surface acceptance insufficiently defined Problem: existing controls do not establish fully objective requirement and/or executed traceable proof for the specific characteristic. | ISO §5.4 (PDF10 / printed4): effective-length outer surface must be free from extraneous matter and harmful processing/surface defects; at least 2.5x magnification applies only if required by risk assessment. | WI p.5 §9 lists irregularities, fish eyes, bubbles and roughness and describes at-least 2.5x as universal; actual ISO condition is risk-based. F-823-i211 rows 10-12 and 15 list appearance defects but no defined defect-classification guide. | §9 (proposed): Visually inspect the specified catheter effective length under normal/corrected vision. Reject foreign material and defects assessed as unacceptable per the approved visual-defect standard. Use magnification >=2.5x only when required by the approved risk assessment or a stricter company specification. Record identified defect and disposition against the work order. | F-823-i12 row 11 "Clear, Good"; F-823-i211 rows 10-12 and 15 define appearance, irregularities, opening defects. | Characteristic: effective-length surface, holes and component appearance; acceptance: approved defect catalogue (ID/rev) and no visible foreign material, damage, flash, blocked holes; actual finding code/location, P/F, method/light/magnification if required. | Required Evidence: Approved visual defect criteria, risk-based magnification decision, signed sample records/defect disposition; not provided. Verification / Control Route: In-process Evidence Reference: Pending / not supplied Closure Criteria: approved applicable control, source-linked execution evidence and authorized disposition. Current Closure Status: OPEN — evidence / technical input / approval; draft wording is not operational closure. |
| GAP-006 ¦ P2 Size, diameter and non-round profile traceability Problem: existing controls do not establish fully objective requirement and/or executed traceable proof for the specific characteristic. | ISO §§5.5.1-5.5.3 (PDF10 / printed4) requires defined nominal outer diameter (mm rounded up to 0.1 mm; +/-1 French except specific exceptions), largest axis for non-round sections, and effective length units; ISO does not provide effective-length tolerance. | WI p.3 §6 requests ID/OD/length "as per specification"; p.4 §8.1.1 gives Redon tube length 500 mm; F-823-i211 rows 13,16-17,20,23 include holes, dimensions and length fields without an approved dimensional tolerance reference. | §§6,10 (proposed): Verify applicable catheter OD / maximum profile axis for flat devices, effective length, ID, overall length and hole dimensions against the SKU-specific approved drawing. Apply ISO nominal size designation rules and its stated OD tolerance unless a documented applicable exception applies. Company drawing shall define all otherwise unspecified limits. Record actual measured values with unit, gauge ID and size reference. | F-823-i12 row 20: Redon 500 mm with holes. F-823-i211 row 13: 1.0 mm hole ID, 51 holes; rows 16-17 width/height/length blanks; row 20 ID/OD/length, row 23 overall length. No approved drawing available. | Both forms: SKU/drawing rev; nominal and tolerance fields from approved drawing; actual OD/maximum axis (mm), effective length (mm/cm), dimension measured, 1.0 mm / 51-hole check only where drawing confirms applicability; measured results with units and gauge IDs. | Required Evidence: Approved size-specific drawings, calibrated measuring equipment and real measurements. OPEN TECHNICAL INPUT REQUIRED for missing drawing-specific limits and configurations. Verification / Control Route: In-process / design specs Evidence Reference: Pending / not supplied Closure Criteria: approved applicable control, source-linked execution evidence and authorized disposition. Current Closure Status: OPEN — evidence / technical input / approval; draft wording is not operational closure. |
| GAP-007 ¦ P1 Connector misconnection safety unverified Problem: existing controls do not establish fully objective requirement and/or executed traceable proof for the specific characteristic. | ISO §5.6 (PDF11 / printed5) requires manufacturers to assess and prevent misconnections based on the general requirements of ISO 80369-1; ISO 20697 does not prescribe one universal connector design. | WI p.2 §5.8 and p.4 §8.1.1 calls for correct connector size/fitting; p.5 §9 uses "strong/no pullout". No connection compatibility/misconnection assessment referenced. | §5.6 (proposed): Only released connector design and mating combinations assessed for misconnection risk may be used. Verify component identity, assembly orientation and correct matching against released drawing. Engineering/QA shall maintain connection safety assessment and qualification evidence; deviations require hold and formal disposition. | F-823-i12 rows 14,17 and 23: good fitting / according to size / strong no pullout. F-823-i211 row18: good fitting/no pullout. | Add connector type and mating component ID, drawing revision, correct connection visual result, misconnection risk-file reference and qualification evidence reference; no universal connector limit invented. | Required Evidence: Approved interface specification and ISO 80369-1-based risk assessment; not supplied. Verification / Control Route: Design/QA qualification + in-process Evidence Reference: Pending / not supplied Closure Criteria: approved applicable control, source-linked execution evidence and authorized disposition. Current Closure Status: OPEN — evidence / technical input / approval; draft wording is not operational closure. |
| GAP-008 ¦ P1 Sterilization qualification and release linkage not verified Problem: existing controls do not establish fully objective requirement and/or executed traceable proof for the specific characteristic. | ISO §5.8 (PDF12 / printed6): supplied-sterile drainage devices must achieve SAL 10^-6 through applicable sterilization controls. | WI p.6 §§12-13 cites WI-7-5-1-1 and review of sterilization cycle and reports. Actual validation, bioburden/sterility assurance and batch release records were not supplied. | §§12-13 (proposed): Sterile product shall not be released before the responsible authorised function verifies the applicable validated sterilization method, approved cycle acceptance, product/load traceability and evidence supporting SAL 10^-6, together with required QA/QC disposition. Link each released batch to the validated process and signed sterilization/release records; reject or hold any missing or failed evidence. | Neither F-823-i12 nor F-823-i211 is a post-sterilization release or batch sterility assurance record. They only document in-process assembly. | Add controlled evidence linkage (not an hourly test): production lot; sterilization lot; procedure/version; cycle record ID; validation scope; QA final disposition reference. Keep sterilization test data in controlled release files. | Required Evidence: Validated sterilization file, approved SAL rationale and batch cycle/release records; WI-7-5-1-1 not supplied. Operational closure unavailable. Verification / Control Route: Sterilization validation / final release Evidence Reference: Pending / not supplied Closure Criteria: approved applicable control, source-linked execution evidence and authorized disposition. Current Closure Status: OPEN — evidence / technical input / approval; draft wording is not operational closure. |
| GAP-009 ¦ P2 Kink criterion and evidence not defined by risk Problem: existing controls do not establish fully objective requirement and/or executed traceable proof for the specific characteristic. | ISO §6.1 (PDF12 / printed6) requires safe catheter placement; the standard gives no universal kink numerical limit. Annex A (PDF17-18 / printed11-12) is informative. | F-823-i211 row 19 mentions "kink resistance" without a method/threshold; WI p.3 §6 checks tube dimensions and appearance but does not specify kink criteria. | §10 (proposed): Where bend/kink performance is identified as relevant by the product risk assessment, Engineering shall approve a clinically justified configuration-specific kink acceptance criterion and verification method, optionally using ISO Annex A. In-process QC shall control related critical extrusion/assembly characteristics and link to qualified kink evidence; do not assign an invented mandrel limit. | F-823-i211 row 19 states kink resistance with no numeric pass/fail condition or evidence. F-823-i12 has no equivalent. | For applicable flat/round tubes: recorded visual status under defined production condition, approved kink-control criterion reference, qualification report ID. Fields for actual measured mandrel diameter (mm) only if an approved test plan requires it. | Required Evidence: Approved clinical/risk justification, kink method/criteria and executed evidence; OPEN TECHNICAL INPUT REQUIRED. Verification / Control Route: Qualification / risk-based periodic Evidence Reference: Pending / not supplied Closure Criteria: approved applicable control, source-linked execution evidence and authorized disposition. Current Closure Status: OPEN — evidence / technical input / approval; draft wording is not operational closure. |
| GAP-010 ¦ P2 Trocar corrosion WI omits specified normative procedure Problem: existing controls do not establish fully objective requirement and/or executed traceable proof for the specific characteristic. | ISO §6.2 (PDF12 / printed6) conditionally requires corrosion testing for exposed metallic parts where corrosion could affect performance; Annex B (PDF19 / printed13) describes saline and thermal conditioning followed by visible-corrosion examination. | WI p.5 §10.1 references Annex B and states trocar must be rust-free; p.4 §8 calls for stainless steel. It does not document test conditions, evidence route, equipment or risk applicability. | §10.1 (proposed): If trocar corrosion risk meets ISO §6.2 applicability, test representative exposed metal to ISO Annex B: 0.15 mol/L NaCl at (22 +/- 5) °C for 5 h; boiling purified water for 30 min; condition at (37 +/- 2) °C for 48 h; dry and inspect visually. Acceptance: no corrosion affecting safe function under approved risk criteria. Retain named test report and specimen/material identification. Routine incoming QC checks for visible corrosion do not replace qualification evidence. | F-823-i12 rows 18-19 track trocar size and sharpness; no corrosion qualification field. F-823-i211 row24 lists trocar dimensions only. | Trocar in-process/incoming: visual corrosion/rust finding P/F, material/lot/drawing; evidence reference to Annex B qualification or approved alternative per risk basis. Do not repeat 48-hour qualification per production check without approved frequency. | Required Evidence: Exposed-metal risk assessment, Annex B test report and traceable incoming verification; no executed reports provided. Verification / Control Route: Qualification / incoming Evidence Reference: Pending / not supplied Closure Criteria: approved applicable control, source-linked execution evidence and authorized disposition. Current Closure Status: OPEN — evidence / technical input / approval; draft wording is not operational closure. |
| GAP-011 ¦ P1 Vacuum deformation lacks defined pressure and method Problem: existing controls do not establish fully objective requirement and/or executed traceable proof for the specific characteristic. | ISO §6.3 (PDF12 / printed6) and Annex C (PDF20 / printed14) require no function-impairing collapse at manufacturer-defined maximum negative pressure; Annex C includes conditioning and at least 60 s pressure exposure. | WI p.5 §10.2 says no deformation/detachment under suction, but no maximum negative pressure, conditioning or objective test record. F-823-i12 row15 is only a manual leak check (different characteristic). | §10.2 (proposed): For devices intended for suction, QA-approved verification shall test the ready-for-use assembly by ISO Annex C. Condition body-insertion portions at 100% RH or in water (37 +/- 2) °C >=15 min and other parts at >=40% RH, (22 +/- 2) °C >=15 min. Immediately apply approved maximum negative pressure for >=60 s. Acceptance: no deformation sufficient to impair function. Record actual pressure, equipment, duration, observation and SKU/lot. | F-823-i12 row15 "No leakage Manual" is not deformation testing. F-823-i211 row19 does not document suction collapse performance. | Add record linkage: negative-pressure deformation qualification ID, maximum approved suction pressure (Pa/kPa), applicable assembly, approved validity. In-process only visual signs/deformation and manufacturing prerequisites, not compulsory per-hour Annex C test. | Required Evidence: Approved max negative pressure, executed Annex C report, equipment calibration and risk-approved verification frequency. OPEN TECHNICAL INPUT REQUIRED for pressure. Verification / Control Route: Qualification / periodic Evidence Reference: Pending / not supplied Closure Criteria: approved applicable control, source-linked execution evidence and authorized disposition. Current Closure Status: OPEN — evidence / technical input / approval; draft wording is not operational closure. |
| GAP-012 ¦ P1 External connector tensile force not quantitatively shown Problem: existing controls do not establish fully objective requirement and/or executed traceable proof for the specific characteristic. | ISO §6.4.1/Table 1 (PDF12 / printed6), Annex D (PDF21 / printed15): for external connections with smallest tube OD 2-4 mm, minimum 5 N; OD >4 mm, minimum 15 N; for OD <2 mm manufacturer sets criterion from risk. Annex D test speed 500 mm/min. | WI p.5 §9 and p.4 §8 use "no pullout"/"strong gluing" but do not identify external connection limits. WI p.6 §10.3 refers to Annex E for "bonding strength", omitting the distinct Annex D connection route. | §10.3a (proposed): For each applicable recommended external connection, perform/maintain ISO Annex D qualification per approved plan with mated parts assembled as instructed, at 500 mm/min and without separation at the applicable Table 1 force (5 N for minimum OD >=2 to <=4 mm; 15 N for >4 mm). For <2 mm, use approved risk-derived criterion. Record OD, applied force, speed, separation and connection identity. | F-823-i12 row23 and F-823-i211 row18 state "Strong, No pull out" or "Not pull out" without numerical loading. | Daily process: correct joining/bond application and visible/physical assembly condition per approved process limit. Qualification link: external connector ID, min tube OD (mm), ISO Table 1 force (N), actual force (N), speed, P/F and Annex D evidence ID. | Required Evidence: Annex D signed tensile results, connector/size matrix, validated bond process and approved monitoring plan; none provided. Verification / Control Route: Qualification + in-process bonding Evidence Reference: Pending / not supplied Closure Criteria: approved applicable control, source-linked execution evidence and authorized disposition. Current Closure Status: OPEN — evidence / technical input / approval; draft wording is not operational closure. |
| GAP-013 ¦ P1 Catheter tubular/junction tensile criteria not determined Problem: existing controls do not establish fully objective requirement and/or executed traceable proof for the specific characteristic. | ISO §6.4.2/Table 2 (PDF13 / printed7), Annex E (PDF22-23 / printed16-17): minimum 10 N for smallest OD 2-4 mm and 20 N for >4 mm; <2 mm risk-derived. Annex E conditioning >=2 h, strain rate 20 mm/min per mm gauge length. | WI p.6 §10.3 mentions Annex E and "standard force" with no N values, tube diameter category, conditioning/gauge length or test record. | §10.3b (proposed): Evaluate every applicable tube portion and junction using Annex E after conditioning insertion parts at (37 +/- 2) °C, 100% RH/water for >=2 h, other components at (22 +/- 5) °C and >=40% RH for >=2 h. Load at 20 mm/min/mm gauge length. For tube OD 2-4 mm require >=10 N; for >4 mm >=20 N; <2 mm criterion is risk-approved. Record OD, test speed, peak force or specified force, location/failure and result. | F-823-i12 row12 "Strong, No pullout"; i211 row18 no pullout, without actual force, test regime or force evidence. | Keep process checks for proper gluing/cure and no visible disengagement. Link applicable design/periodic Annex E force test by SKU, tube category, junction ID, OD, gauge length, actual N, test date and evidence ID. | Required Evidence: Annex E report, bonding process qualification and change-control scope; not supplied. Verification / Control Route: Qualification / periodic + in-process Evidence Reference: Pending / not supplied Closure Criteria: approved applicable control, source-linked execution evidence and authorized disposition. Current Closure Status: OPEN — evidence / technical input / approval; draft wording is not operational closure. |
| GAP-014 ¦ P1 Collection device impact/vacuum retention verification missing Problem: existing controls do not establish fully objective requirement and/or executed traceable proof for the specific characteristic. | ISO §6.5 (PDF13 / printed7) and Annex F (PDF24-25 / printed18-19): collection device must not leak after drop; suction source vacuum loss must not exceed 2%; 700 mm drop at (22 +/- 5) °C, vacuum assessed >=60 s after impact. | WI p.2 §5.3 asks bellows free from holes/cracks and leakage. Neither WI nor in-process records cites Annex F or drop performance. | §10.4a (proposed): Verify collection device impact performance under an approved ISO Annex F plan: assemble as clinically used, fill collection device to effective capacity, occlude and drop 700 mm onto specified rigid surface at (22 +/- 5) °C; acceptance no leakage. For any suction-source configuration, record initial negative pressure and pressure at >=60 s post-impact; acceptance vacuum loss <=2%. Retain configuration-specific objective evidence. | F-823-i12 rows 11,15,16 inspect bellows appearance, manual leak, nominal capacity only, not impact; flat assembly form is not a collection-device test. | Link collection-device model, capacity, applicable suction-source variant, Annex F qualification report ID, drop conditions, leakage P/F and percent vacuum loss. Routine visual bellow control retained. | Required Evidence: Impact qualification test reports and controlled design/lot change coverage. Not supplied. Verification / Control Route: Qualification / periodic Evidence Reference: Pending / not supplied Closure Criteria: approved applicable control, source-linked execution evidence and authorized disposition. Current Closure Status: OPEN — evidence / technical input / approval; draft wording is not operational closure. |
| GAP-015 ¦ P2 Conditional claimed flow performance not traceable Problem: existing controls do not establish fully objective requirement and/or executed traceable proof for the specific characteristic. | ISO §6.6 (PDF13 / printed7) and Annex G (PDF26-27 / printed20-21): ONLY where manufacturer claims flow, each lumen must meet >=80% of claimed flow for OD <1.0 mm, >=90% for OD >=1.0 mm; Annex G uses three determinations and reports average mL/min. | WI p.3 §6.2 requires holes free from blockage; p.4-5 checks punched holes. No product flow claim provided, so applicability cannot be determined. | §10.6 (proposed): Engineering/QA shall document whether the configuration has a manufacturer flow-rate claim. If yes, verify each lumen according to ISO Annex G, with the prescribed 1000 +/-5 mm head, 22 +/-5 °C liquid and >=30 s collection for three readings. Accept mean rate >= applicable 80%/90% of approved claim. Record claimed rate, lumen, three rates, average and lot/configuration. If no claim, retain documented N/A rationale; maintain unobstructed-hole in-process checks. | F-823-i12 row22 "without ... blockage"; i211 rows 13-15 require open perforations; no flow rate field or claim evidence. | In-process: holes open, number/size per released drawing with actual findings. External verification index: claimed flow mL/min, diameter category, result of three determinations, average and report ID if claimed. | Required Evidence: Manufacturer claims/IFU, Annex G test evidence or approved N/A assessment; OPEN TECHNICAL INPUT REQUIRED on flow claim. Verification / Control Route: Conditional qualification Evidence Reference: Pending / not supplied Closure Criteria: approved applicable control, source-linked execution evidence and authorized disposition. Current Closure Status: OPEN — evidence / technical input / approval; draft wording is not operational closure. |
| GAP-016 ¦ P1 Manual bellows leak check does not establish Annex M vacuum leakage conformity Problem: existing controls do not establish fully objective requirement and/or executed traceable proof for the specific characteristic. | ISO §6.11 (PDF14 / printed8), Annex M (PDF41-42 / printed35-36) require no air ingress at manufacturer-defined maximum negative pressure; Annex M water test at (22 +/- 5) °C includes 120 s stabilization plus 120 s maintained observation. | WI p.2 §5.3 and p.4 §8.1.1 demand bellows leak-free. F-823-i12 row15 "No leakage Manual" has no stated pressure, detection method, duration, result or Annex M evidence. | §10.7 (proposed): Qualify the complete ready-to-use suction assembly under Annex M at approved maximum negative pressure using de-aerated distilled/deionized water at (22 +/- 5) °C. Vent entrained air, occlude appropriately, allow 120 s stabilization and maintain pressure for a further 120 s; acceptance no observed air-bubble ingress. Record actual test pressure, duration, location of any ingress and report ID. Keep routine manual leak checks only under an approved validated process check. | F-823-i12 row15 "Leak test for Bellow pump / No leakage Manual"; no corresponding ISO pressure/time parameters. | Replace ambiguous manual check instruction with approved in-process leak-check procedure ID, actual method / setpoint / hold time if defined, P/F and operator; add independent Annex M evidence reference and pressure/risk approval link. | Required Evidence: Approved maximum pressure, Annex M executed qualification, validated in-process manual test parameters and records; OPEN TECHNICAL INPUT REQUIRED for routine pressure/time. Verification / Control Route: In-process + qualification Evidence Reference: Pending / not supplied Closure Criteria: approved applicable control, source-linked execution evidence and authorized disposition. Current Closure Status: OPEN — evidence / technical input / approval; draft wording is not operational closure. |
| GAP-017 ¦ P2 Capacity, pre-evacuated vacuum claim and label criteria unlinked Problem: existing controls do not establish fully objective requirement and/or executed traceable proof for the specific characteristic. | ISO §7.2(d)-(f) (PDF15 / printed9) requires applicable effective capacity in mL, vacuum-stability date for pre-evacuated sources (>=80% of initial labelled negative pressure), and applicable maximum negative pressure in Pa on packaging. | WI p.4 §8.1.1 lists 200/400/600 mL as possible nominal capacities; F-823-i12 row16 repeats sizes but has no actual measured capacity, permitted tolerance or vacuum-stability test/label record. | §§8,11 (proposed): Match bellows nominal size and effective collection capacity to approved SKU and labelled claim. Verify volume using approved water-based method, recorded volume (mL) and manufacturer-controlled tolerance; absence of a controlled tolerance shall be OPEN TECHNICAL INPUT REQUIRED. For a pre-evacuated suction source, QA shall maintain approved vacuum-stability evidence and label expiry/use-until decision demonstrating >=80% remaining labelled initial vacuum at the specified date. Verify applicable maximum suction pressure information. | F-823-i12 row16 "Capacity (ml): 200ml, 400ml, 600ml"; no actual result or product-specific tolerance; flat form no collection capacity row. | Bellow capacity: SKU nominal capacity (mL), approved acceptance range, actual water volume (mL), result and gauge. Link labelled initial vacuum/stability and maximum suction pressure approval/evidence; retain N/A if source not pre-evacuated. | Required Evidence: Approved bellows spec, volume tolerance, stable-vacuum qualification and approved artwork; no data provided. Verification / Control Route: In-process + qualification/label Evidence Reference: Pending / not supplied Closure Criteria: approved applicable control, source-linked execution evidence and authorized disposition. Current Closure Status: OPEN — evidence / technical input / approval; draft wording is not operational closure. |
| GAP-018 ¦ P2 Packaging label/IFU content review evidence incomplete Problem: existing controls do not establish fully objective requirement and/or executed traceable proof for the specific characteristic. | ISO §§7.1-7.3 (PDF15-16 / printed9-10) require durable/legible device marking; packaging and supplied instructions must include prescribed information where relevant, such as handling, single-use status, collection capacity, negative pressure and appropriate warnings/revision identity. | WI p.6 §11 requires label product name, code, lot, size, expiry and general print/seal checks, but does not explicitly link the full applicable ISO §§7.1-7.3 labelling/IFU verification set and approved artwork. | §11 (proposed): Packaging operators shall use only QA-approved artwork, IFU and packaging specifications for SKU/revision. QC shall verify identity, lot, quantity, required sterile/single-use statements, capacity/max pressure/detectability and other §7.2 fields wherever applicable; ensure IFU has pertinent use/connection warnings, revision and contraindications where relevant. Check marking legibility/durability by validated method and seal formation per approved packaging process. Do not substitute a generic checklist for released artwork. | F-823-i12 and F-823-i211 are assembly-control forms; neither is packaging/final label review. WI p.7 §14 cites F-823-i4-2, not supplied. | Controlled reference only on in-process record: packaging artwork ID/rev and downstream packaging check form ID. Perform and document detailed label/seal checks in existing controlled packaging/final records after verification. | Required Evidence: Approved artworks/IFUs, packaging controls/validation and completed lot packaging records F-823-i4-2; not supplied. Verification / Control Route: Packaging / final release Evidence Reference: Pending / not supplied Closure Criteria: approved applicable control, source-linked execution evidence and authorized disposition. Current Closure Status: OPEN — evidence / technical input / approval; draft wording is not operational closure. |
| GAP-019 ¦ P2 In-process characteristics lack objective result capture / approved technical limits Problem: existing controls do not establish fully objective requirement and/or executed traceable proof for the specific characteristic. | Company process-control issue informed by WI p.4-5 §8.1.1; ISO §5.4 and §5.5 only support relevant surface/dimension expectations. ISO does NOT prescribe 51 perforations, 1.0 mm holes, 500 mm Redon length or 200/400/600 mL SKU tolerances. | F-823-i12 rows 11-27 use 12 hour columns plus general remarks, with terms "good", "strong" and "as per requirement". F-823-i211 rows 13,16-17,20,23 include design-specific dimensions and 1.0 mm/51-hole control but do not specify approved drawing tolerance or systematically capture actual measurements at each checkpoint. | §§8.1,8.2,10.8 (proposed): At each scheduled in-process checkpoint, verify only characteristics applicable to the current SKU/operation using released component drawings and a QA-approved process control plan. Quantitative criteria, gauges, samples and acceptance disposition shall come from those controlled records. Record actual dimension/measurement and units where relevant; qualitative observations use defined defect codes. Record time, shift, traceable lot, operator/QC initials, nonconformance reference and supervisor review. Failed/pending checks require segregation and documented disposition before continuation. | F-823-i12 hourly series G10:R10 = 7,8,9,10,11,12,1,2,3,4,5,6; 17 traits rows 11-27. F-823-i211 E9:P9 = 7,9,11,1,3,5,7,9,11,1,3,5; 13 traits rows 10-24. Header has equipment free-text and blank result grid. | Two proposed editable in-process forms (not final inspection): every checkpoint includes source time pattern, component/SKU/drawing, acceptance reference, actual or defect code, equipment, Pass/Fail/Pending, corrective action link, QC signature and manager review. Preserve 17 and 13 distinct source characteristics, conditional features and original nominal values until controlled revision. | Required Evidence: Approved TF-23-PS/TF-23-PDBM, drawings, sampling/control plan, instrument calibration and executed signed in-process records; none supplied. OPEN TECHNICAL INPUT REQUIRED for unspecified tolerances and tests. Verification / Control Route: Routine in-process / Document Control Evidence Reference: Pending / not supplied Closure Criteria: approved applicable control, source-linked execution evidence and authorized disposition. Current Closure Status: OPEN — evidence / technical input / approval; draft wording is not operational closure. |
| GAP-020 ¦ P2 MRI and optional feature applicability not formally recorded Problem: existing controls do not establish fully objective requirement and/or executed traceable proof for the specific characteristic. | ISO §5.7 (PDF12 / printed6) addresses MRI hazards if applicable; §§5.5.4, 6.7-6.10 and relevant Annexes H-L apply only to retained or balloon configurations meeting their prerequisites (PDF11-14 / printed5-8). | WI pp.1-7 describes standard wound drainage and some X-ray line/trocar variants, but does not establish whether MRI-related usage, retention features or balloon components are included in the current product family. | §5.3 (proposed): QA/Engineering shall issue a per-SKU applicability matrix for MRI use, retention mechanisms, compliant/noncompliant balloons, aspiration vacuum, flow claims and available accessory variants. For every feature present, adopt applicable ISO test methods and criteria with evidence IDs; for absent features mark N/A only with approved drawing/label/risk justification. MRI hazards, when applicable, shall be evaluated by appropriate validated methods. | F-823-i12 and i211 do not have balloon or MRI checkpoints; this is not proof they are present or absent. Some optional ISO clauses cannot yet be dispositioned. | Form applicability header: SKU, configuration/feature flags, approved feature matrix reference. No automatic balloon tests added to routine in-process forms unless documented applicable. | Required Evidence: Approved BOM/drawing, intended use and risk assessment with signed applicability verdict; not supplied. Verification / Control Route: Design applicability / QA Evidence Reference: Pending / not supplied Closure Criteria: approved applicable control, source-linked execution evidence and authorized disposition. Current Closure Status: OPEN — evidence / technical input / approval; draft wording is not operational closure. |

# 6. Required tests and controls | do not misclassify in-process as final inspection

ROUTINE IN-PROCESS: scheduled visual, dimensional, assembly, component completeness, bellows leak check per approved procedure, tube perforation and drawing-specific configuration checks. Existing timing grids (i12 hourly; i211 every two hours in the source) are preserved as observation schedules, subject to an approved control plan.

QUALIFICATION / CHANGE-BASED / PERIODIC PER APPROVED PLAN: Annex B corrosion if required; C negative-pressure deformation; D external connector tensile; E catheter/junction tensile; F collection device impact; M air leakage; A kink only if appropriate, G flow only if claimed. Standards do not automatically require each test per lot.

DESIGN / QUALITY SYSTEM EVIDENCE: ISO 10993-1 biological evaluation, ISO 80369-1 misconnection risk, MRI / detectability / retention and balloon feature assessments as applicable, process/sterilization validation and label/IFU verification.

INCOMING / PACKAGING / FINAL RELEASE: existing WI cites OBM incoming, packaging F-823-i4-2 and final release F-8-2-4; these source forms and completed records were not included and are NOT reconstructed as if reviewed.

# 7. Manager corrective action plan

| Gap ID | Priority | Required action / decision | Responsible function | Required evidence | Closure status |
| --- | --- | --- | --- | --- | --- |
| GAP-001 | P2 | Controlled document reference conflict; approve proposed revised text and controlled applicability/verification route | Document Control | Approved document index, WI/form cross-reference, master list entry, documented revision decisions. No approved master list supplied. | OPEN — evidence / approval |
| GAP-002 | P1 | Intended performance and risk-management traceability; approve proposed revised text and controlled applicability/verification route | Design/QA qualification | Approved risk-management file and report, design verification matrix, product configurations and design claim evidence; none provided. | OPEN — evidence / approval |
| GAP-003 | P1 | Biological evaluation evidence not established; approve proposed revised text and controlled applicability/verification route | Design/QA qualification | Approved biological evaluation report/plan, material specs, supplier traceability and change assessments; not supplied. | OPEN — evidence / approval |
| GAP-004 | P1 | Detectability / X-ray claim verification unclear; approve proposed revised text and controlled applicability/verification route | In-process plus qualification | Detectability risk decision, approved drawing and performance qualification evidence, not supplied. | OPEN — evidence / approval |
| GAP-005 | P2 | Surface acceptance insufficiently defined; approve proposed revised text and controlled applicability/verification route | In-process | Approved visual defect criteria, risk-based magnification decision, signed sample records/defect disposition; not provided. | OPEN — evidence / approval |
| GAP-006 | P2 | Size, diameter and non-round profile traceability; approve proposed revised text and controlled applicability/verification route | In-process / design specs | Approved size-specific drawings, calibrated measuring equipment and real measurements. OPEN TECHNICAL INPUT REQUIRED for missing drawing-specific limits and configurations. | OPEN — evidence / approval |
| GAP-007 | P1 | Connector misconnection safety unverified; approve proposed revised text and controlled applicability/verification route | Design/QA qualification + in-process | Approved interface specification and ISO 80369-1-based risk assessment; not supplied. | OPEN — evidence / approval |
| GAP-008 | P1 | Sterilization qualification and release linkage not verified; approve proposed revised text and controlled applicability/verification route | Sterilization validation / final release | Validated sterilization file, approved SAL rationale and batch cycle/release records; WI-7-5-1-1 not supplied. Operational closure unavailable. | OPEN — evidence / approval |
| GAP-009 | P2 | Kink criterion and evidence not defined by risk; approve proposed revised text and controlled applicability/verification route | Qualification / risk-based periodic | Approved clinical/risk justification, kink method/criteria and executed evidence; OPEN TECHNICAL INPUT REQUIRED. | OPEN — evidence / approval |
| GAP-010 | P2 | Trocar corrosion WI omits specified normative procedure; approve proposed revised text and controlled applicability/verification route | Qualification / incoming | Exposed-metal risk assessment, Annex B test report and traceable incoming verification; no executed reports provided. | OPEN — evidence / approval |
| GAP-011 | P1 | Vacuum deformation lacks defined pressure and method; approve proposed revised text and controlled applicability/verification route | Qualification / periodic | Approved max negative pressure, executed Annex C report, equipment calibration and risk-approved verification frequency. OPEN TECHNICAL INPUT REQUIRED for pressure. | OPEN — evidence / approval |
| GAP-012 | P1 | External connector tensile force not quantitatively shown; approve proposed revised text and controlled applicability/verification route | Qualification + in-process bonding | Annex D signed tensile results, connector/size matrix, validated bond process and approved monitoring plan; none provided. | OPEN — evidence / approval |
| GAP-013 | P1 | Catheter tubular/junction tensile criteria not determined; approve proposed revised text and controlled applicability/verification route | Qualification / periodic + in-process | Annex E report, bonding process qualification and change-control scope; not supplied. | OPEN — evidence / approval |
| GAP-014 | P1 | Collection device impact/vacuum retention verification missing; approve proposed revised text and controlled applicability/verification route | Qualification / periodic | Impact qualification test reports and controlled design/lot change coverage. Not supplied. | OPEN — evidence / approval |
| GAP-015 | P2 | Conditional claimed flow performance not traceable; approve proposed revised text and controlled applicability/verification route | Conditional qualification | Manufacturer claims/IFU, Annex G test evidence or approved N/A assessment; OPEN TECHNICAL INPUT REQUIRED on flow claim. | OPEN — evidence / approval |
| GAP-016 | P1 | Manual bellows leak check does not establish Annex M vacuum leakage conformity; approve proposed revised text and controlled applicability/verification route | In-process + qualification | Approved maximum pressure, Annex M executed qualification, validated in-process manual test parameters and records; OPEN TECHNICAL INPUT REQUIRED for routine pressure/time. | OPEN — evidence / approval |
| GAP-017 | P2 | Capacity, pre-evacuated vacuum claim and label criteria unlinked; approve proposed revised text and controlled applicability/verification route | In-process + qualification/label | Approved bellows spec, volume tolerance, stable-vacuum qualification and approved artwork; no data provided. | OPEN — evidence / approval |
| GAP-018 | P2 | Packaging label/IFU content review evidence incomplete; approve proposed revised text and controlled applicability/verification route | Packaging / final release | Approved artworks/IFUs, packaging controls/validation and completed lot packaging records F-823-i4-2; not supplied. | OPEN — evidence / approval |
| GAP-019 | P2 | In-process characteristics lack objective result capture / approved technical limits; approve proposed revised text and controlled applicability/verification route | Routine in-process / Document Control | Approved TF-23-PS/TF-23-PDBM, drawings, sampling/control plan, instrument calibration and executed signed in-process records; none supplied. OPEN TECHNICAL INPUT REQUIRED for unspecified tolerances and tests. | OPEN — evidence / approval |
| GAP-020 | P2 | MRI and optional feature applicability not formally recorded; approve proposed revised text and controlled applicability/verification route | Design applicability / QA | Approved BOM/drawing, intended use and risk assessment with signed applicability verdict; not supplied. | OPEN — evidence / approval |

# 8. Closure methodology and final conclusion

Documentary draft coverage: proposed text maps the listed ISO clauses to explicit controls, but unapproved drafts and unknown SKU-specific criteria prevent a defensible 100% coverage conclusion. Operational conformity: NOT VERIFIED because no completed sampling, signed test/qualification data, current master drawings or signed validated release evidence were supplied. Exact applicable clauses for MRI/balloon/retention/flow require manufacturer risk and intended-use decisions.

QC Manager approvals requested: (1) confirm current controlled WI ID and both form reference revisions, (2) approve applicability-by-SKU matrix, (3) assign Engineering/QA to release quantitative limits and max suction pressure, (4) approve risk-based qualification and in-process frequencies, (5) approve controlled documentation after objective evidence and test methods verified. All deliverables remain draft.