# REVISED IN-PROCESS CONTROL FORMS

Two complete editable working templates (F-823-i12 and F-823-i211 source baselines) • NOT final inspection forms

| PRODUCT Wound Drainage Set / flat perforated drain variants STANDARD ISO 20697:2018(E), corrected version 2018-09 | SOURCE WI WI-8-2-2-23 Rev.4, 25 Dec 2022 STATUS Controlled draft; no approval implied |
| --- | --- |

Evidence boundary: one official ISO PDF, one signed WI scan and two blank in-process XLSX forms were supplied. No executed records, master drawings, test reports, sterilization validation, risk file, or approved technical specifications were provided. A gap in supplied evidence is NOT evidence of a failed physical test.

# Purpose, usage and record status

This deliverable reconstructs ONLY the two supplied in-process production control records. The source form numbers, 17+13 original checkpoints and hour schedules are retained. Draft proposed fields are editable blank entries: replace controlled drawings/limits only after approval. Do not use as current effective controlled forms. Qualification tests for ISO 20697 belong in controlled separate reports, with evidence links below.

# A. Wound Drainage Set

Controlled source: F-823-i12 Rev.4, 23 Nov 2025; proposed future revision: TBD; current source grid checkpoint labels: 7, 8, 9, 10, 11, 12, 1, 2, 3, 4, 5, 6. The time labels are transcribed from the supplied source form; shift/AM-PM interpretation must be established by the approved site schedule.

| Record metadata | Editable required entry | Record metadata | Editable required entry |
| --- | --- | --- | --- |
| Work order / lot | ________________ | SKU / size / variant | ________________ |
| Material or component lot | ________________ | Drawing / BOM + rev | ________________ |
| WI ID/revision | ________________ | Form ID / draft rev | F-823-i12 / TBD |
| Date / shift | ________________ | Approved control-plan ID/rev | ________________ |
| Machine / assembly station | ________________ | QC inspector | ________________ |
| Instruments / calibration IDs | ________________ | Supervisor reviewer | ________________ |

## A. Master characteristic and technical control register

| ID | Characteristic | Baseline requirement (must confirm SKU/drawing) | Test method / record | Source row | Approved limit / drawing rev (editable) | Measured/observed actual | Disposition |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | Tube / bellows pump clarity | Clear / good | Visual | 11 | ________________ | ________________ | P / F / Pending |
| 2 | Tube gluing / insertion | Strong; no pullout | Visual / physical (source WI) | 12 | ________________ | ________________ | P / F / Pending |
| 3 | Tube with slide clamp | Available | Visual | 13 | ________________ | ________________ | P / F / Pending |
| 4 | Fittings status | Good fitting | Visual | 14 | ________________ | ________________ | P / F / Pending |
| 5 | Leak test for bellows pump | No leakage; manual | Manual per approved WI / pending details | 15 | ________________ | ________________ | P / F / Pending |
| 6 | Capacity (mL) | 200 / 400 / 600 mL nominal variants | Water measurement; tolerance pending | 16 | ________________ | ________________ | P / F / Pending |
| 7 | Connector type | According to size | Visual against drawing | 17 | ________________ | ________________ | P / F / Pending |
| 8 | Trocar | According to size | Visual / measure against drawing | 18 | ________________ | ________________ | P / F / Pending |
| 9 | Trocar sharpness | Sharp; no damage | Visual / physical | 19 | ________________ | ________________ | P / F / Pending |
| 10 | Redon drain tube length | 500 mm with multiple holes; only if this SKU | Measure; drawing tolerance pending | 20 | ________________ | ________________ | P / F / Pending |
| 11 | Bellows graduation scale | Clear marking | Visual | 21 | ________________ | ________________ | P / F / Pending |
| 12 | Punching-hole quality | No fibres, blockage, roughness | Visual | 22 | ________________ | ________________ | P / F / Pending |
| 13 | Connector to bellows | Strong; no pullout | Visual / physical | 23 | ________________ | ________________ | P / F / Pending |
| 14 | Correct specification | Per approved SKU specification | Visual / measure as applicable | 24 | ________________ | ________________ | P / F / Pending |
| 15 | Hole at bellows neck | Clear; no fibres | Visual | 25 | ________________ | ________________ | P / F / Pending |
| 16 | Rubber on hole | Yes / No per approved configuration | Visual | 26 | ________________ | ________________ | P / F / Pending |
| 17 | Hanger pieces | Available if specified | Visual | 27 | ________________ | ________________ | P / F / Pending |

## Scheduled checkpoint completion grid

For each corresponding observation time, enter P=pass, F=fail, H=held/pending or N/A with approved justification; actual numeric values and findings MUST also be captured in the measurement/exception record below or the traceable lab worksheet. This grid alone is not a numeric measurement record.

| ID / time | 7 | 8 | 9 | 10 | 11 | 12 | 1 | 2 | 3 | 4 | 5 | 6 | Remarks / NCR |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 Tube / bellows pump clar | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ___________ |
| 2 Tube gluing / insertion | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ___________ |
| 3 Tube with slide clamp | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ___________ |
| 4 Fittings status | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ___________ |
| 5 Leak test for bellows pu | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ___________ |
| 6 Capacity (mL) | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ___________ |
| 7 Connector type | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ___________ |
| 8 Trocar | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ___________ |
| 9 Trocar sharpness | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ___________ |
| 10 Redon drain tube length | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ___________ |
| 11 Bellows graduation scale | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ___________ |
| 12 Punching-hole quality | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ___________ |
| 13 Connector to bellows | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ___________ |
| 14 Correct specification | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ___________ |
| 15 Hole at bellows neck | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ___________ |
| 16 Rubber on hole | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ___________ |
| 17 Hanger pieces | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ___________ |

## Actual measurements, deviations and checkpoint signatures

| Date/time | ID | Sample ID / quantity | Actual reading or finding / units | Acceptance range + drawing | Instrument ID | Result | NCR / disposition |
| --- | --- | --- | --- | --- | --- | --- | --- |
| ________ | ____ | ________ | __________________ | ________________ | ________ | P/F/H | ____________ |
| ________ | ____ | ________ | __________________ | ________________ | ________ | P/F/H | ____________ |
| ________ | ____ | ________ | __________________ | ________________ | ________ | P/F/H | ____________ |
| ________ | ____ | ________ | __________________ | ________________ | ________ | P/F/H | ____________ |
| ________ | ____ | ________ | __________________ | ________________ | ________ | P/F/H | ____________ |
| ________ | ____ | ________ | __________________ | ________________ | ________ | P/F/H | ____________ |
| ________ | ____ | ________ | __________________ | ________________ | ________ | P/F/H | ____________ |
| ________ | ____ | ________ | __________________ | ________________ | ________ | P/F/H | ____________ |

QC inspected by / signature / date/time: __________________________________  Reviewed by QC Supervisor/QCM per effective approvals: __________________________________

Nonconformity: failed/held characteristic must have affected product range, physical/system segregation reference, authorised disposition and re-check outcome. Sample size, equipment and limits from approved control plan; do not invent values.

## Traceability to ISO / separate evidence (not additional routine testing)

| Gap / ISO link | External control | Applicable evidence scope | Validated record ID |
| --- | --- | --- | --- |
| GAP-011 / REQ-13 | Annex C suction deformation | Approved qualification record ID / applicable SKU, pressure (pending) | _______________ |
| GAP-012 / REQ-14 | Annex D connection tensile | Approved force qualification (5/15 N by OD) | _______________ |
| GAP-013 / REQ-15 | Annex E tube/junction tensile | Approved force qualification (10/20 N by OD) | _______________ |
| GAP-014 / REQ-16 | Annex F impact | Collection / suction qualification | _______________ |
| GAP-016 / REQ-22 | Annex M air ingress | Validated ready-for-use pressure/leakage qualification | _______________ |
| GAP-008 / REQ-10 | Sterilization | Sterilization lot/validation/release reference in final controlled record | _______________ |

# B. Silicone Flat Perforated Drain Tube Assembly

Controlled source: F-823-i211 Rev.4, 06 Dec 2025; proposed future revision: TBD; current source grid checkpoint labels: 7, 9, 11, 1, 3, 5, 7, 9, 11, 1, 3, 5. The time labels are transcribed from the supplied source form; shift/AM-PM interpretation must be established by the approved site schedule.

| Record metadata | Editable required entry | Record metadata | Editable required entry |
| --- | --- | --- | --- |
| Work order / lot | ________________ | SKU / size / variant | ________________ |
| Material or component lot | ________________ | Drawing / BOM + rev | ________________ |
| WI ID/revision | ________________ | Form ID / draft rev | F-823-i211 / TBD |
| Date / shift | ________________ | Approved control-plan ID/rev | ________________ |
| Machine / assembly station | ________________ | QC inspector | ________________ |
| Instruments / calibration IDs | ________________ | Supervisor reviewer | ________________ |

## B. Master characteristic and technical control register

| ID | Characteristic | Baseline requirement (must confirm SKU/drawing) | Test method / record | Source row | Approved limit / drawing rev (editable) | Measured/observed actual | Disposition |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | Appearance | Clear and good appearance | Visual | 10 | ________________ | ________________ | P / F / Pending |
| 2 | Freedom from irregularities | No dust, black spots, fish eyes | Visual | 11 | ________________ | ________________ | P / F / Pending |
| 3 | Flat tube status | White silicone flat perforated; smooth; no cracking, damage, obstruction | Visual | 12 | ________________ | ________________ | P / F / Pending |
| 4 | Flat tube hole drilling | 1.0 mm hole ID and 51 holes, only for specified configuration; no sharp edges | Measure / count; drawing tolerance pending | 13 | ________________ | ________________ | P / F / Pending |
| 5 | Channel status | Per approved size-specific drawing | Visual / measure | 14 | ________________ | ________________ | P / F / Pending |
| 6 | Flat tube perforations | Open; no flash or blockage | Visual | 15 | ________________ | ________________ | P / F / Pending |
| 7 | Flat tube dimensions | Width / height / length per drawing | Measure | 16-17 | ________________ | ________________ | P / F / Pending |
| 8 | Flat-to-round tube/adaptor assembly | No flash; fit secure; no pullout | Visual / physical | 18 | ________________ | ________________ | P / F / Pending |
| 9 | Round tube status | Transparent silicone, visible X-ray line, kink control, no holes/leak | Visual; other performance via qualification | 19 | ________________ | ________________ | P / F / Pending |
| 10 | Round tube dimensions | ID/OD/length per approved drawing | Measure | 20 | ________________ | ________________ | P / F / Pending |
| 11 | White clamp | Smooth, no flash, cracks/breakage, locks | Visual / functional | 22 | ________________ | ________________ | P / F / Pending |
| 12 | Total length | Per approved drawing in cm | Measure | 23 | ________________ | ________________ | P / F / Pending |
| 13 | Trocar | Applicable size, OD and length | Measure / visual | 24 | ________________ | ________________ | P / F / Pending |

## Scheduled checkpoint completion grid

For each corresponding observation time, enter P=pass, F=fail, H=held/pending or N/A with approved justification; actual numeric values and findings MUST also be captured in the measurement/exception record below or the traceable lab worksheet. This grid alone is not a numeric measurement record.

| ID / time | 7 | 9 | 11 | 1 | 3 | 5 | 7 | 9 | 11 | 1 | 3 | 5 | Remarks / NCR |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 Appearance | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ___________ |
| 2 Freedom from irregularit | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ___________ |
| 3 Flat tube status | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ___________ |
| 4 Flat tube hole drilling | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ___________ |
| 5 Channel status | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ___________ |
| 6 Flat tube perforations | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ___________ |
| 7 Flat tube dimensions | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ___________ |
| 8 Flat-to-round tube/adapt | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ___________ |
| 9 Round tube status | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ___________ |
| 10 Round tube dimensions | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ___________ |
| 11 White clamp | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ___________ |
| 12 Total length | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ___________ |
| 13 Trocar | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ____ | ___________ |

## Actual measurements, deviations and checkpoint signatures

| Date/time | ID | Sample ID / quantity | Actual reading or finding / units | Acceptance range + drawing | Instrument ID | Result | NCR / disposition |
| --- | --- | --- | --- | --- | --- | --- | --- |
| ________ | ____ | ________ | __________________ | ________________ | ________ | P/F/H | ____________ |
| ________ | ____ | ________ | __________________ | ________________ | ________ | P/F/H | ____________ |
| ________ | ____ | ________ | __________________ | ________________ | ________ | P/F/H | ____________ |
| ________ | ____ | ________ | __________________ | ________________ | ________ | P/F/H | ____________ |
| ________ | ____ | ________ | __________________ | ________________ | ________ | P/F/H | ____________ |
| ________ | ____ | ________ | __________________ | ________________ | ________ | P/F/H | ____________ |
| ________ | ____ | ________ | __________________ | ________________ | ________ | P/F/H | ____________ |
| ________ | ____ | ________ | __________________ | ________________ | ________ | P/F/H | ____________ |

QC inspected by / signature / date/time: __________________________________  Reviewed by QC Supervisor/QCM per effective approvals: __________________________________

Nonconformity: failed/held characteristic must have affected product range, physical/system segregation reference, authorised disposition and re-check outcome. Sample size, equipment and limits from approved control plan; do not invent values.

## Traceability to ISO / separate evidence (not additional routine testing)

| Gap / ISO link | External control | Applicable evidence scope | Validated record ID |
| --- | --- | --- | --- |
| GAP-011 / REQ-13 | Annex C suction deformation | Approved qualification record ID / applicable SKU, pressure (pending) | _______________ |
| GAP-012 / REQ-14 | Annex D connection tensile | Approved force qualification (5/15 N by OD) | _______________ |
| GAP-013 / REQ-15 | Annex E tube/junction tensile | Approved force qualification (10/20 N by OD) | _______________ |
| GAP-014 / REQ-16 | Annex F impact | Collection / suction qualification | _______________ |
| GAP-016 / REQ-22 | Annex M air ingress | Validated ready-for-use pressure/leakage qualification | _______________ |
| GAP-008 / REQ-10 | Sterilization | Sterilization lot/validation/release reference in final controlled record | _______________ |