QUALITY SYSTEMS / CONTROLLED DRAFT

# ISO 20697 Closure Register

Applicable requirements, documentary status, executed evidence and open technical actions

Wound Drainage Set   |   ISO 20697:2018(E), first edition 2018-06, corrected 2018-09

Status: DRAFT FOR QC/QA/ENGINEERING REVIEW   •   No actual product conformity or ISO closure is asserted.

Source documents: WI-8-2-2-23; F-823-i12; F-823-i108; ISO 20697:2018(E). Evidence supplied consists of unexecuted document templates, not completed production/test records.

# Closure rules and authorization

Verified Closed requires an effective approved controlled requirement AND applicable reviewed executed objective evidence. A template, a draft WI, an unsigned certification or an ISO citation is not evidence of conformity. All gap rows below are open until approved status is changed against reviewed records. No “100% conformance” finding is possible from these source documents.

# Clause-by-Clause Applicability and Traceability

| REQ ID | ISO clause | ISO says (summary) | Gap/change mapping | Documentary coverage | Verified operational evidence | Closure evidence / outstanding action | Status |
| --- | --- | --- | --- | --- | --- | --- | --- |
| REQ-001 | 4; 5.1 | Intended drainage function / risk process | GAP-002 | Draft proposed / NOT approved | Not provided | Risk/design verification evidence pending; effective approved record and actual traceable test evidence as applicable | Open — evidence/technical decision |
| REQ-002 | 5.2 | Biocompatibility | GAP-003 | Draft proposed / NOT approved | Not provided | Not evidenced; effective approved record and actual traceable test evidence as applicable | Open — evidence/technical decision |
| REQ-003 | 5.3 | Detectability if risk-required | GAP-004 | Draft proposed / NOT approved | Not provided | Conditional; risk decision missing; effective approved record and actual traceable test evidence as applicable | Open — evidence/technical decision |
| REQ-004 | 5.4 | Surface finish | GAP-005 | Draft proposed / NOT approved | Not provided | Qualitatively addressed; risk magnification linkage pending; effective approved record and actual traceable test evidence as applicable | Open — evidence/technical decision |
| REQ-005 | 5.5.1–5.5.3 | OD and length designation | GAP-006 / GAP-021 | Draft proposed / NOT approved | Not provided | Covered partly; drawing/tolerances pending; effective approved record and actual traceable test evidence as applicable | Open — evidence/technical decision |
| REQ-006 | 5.5.4 | Balloon inflation volume | N/A pending design proof | N/A — provisional | Not provided | Controlled product-config/BOM proof of no catheter balloon; QA applicability authorization | Pending applicability proof |
| REQ-007 | 5.6 | Connector misconnection | GAP-007 | Draft proposed / NOT approved | Not provided | Risk/design proof pending; effective approved record and actual traceable test evidence as applicable | Open — evidence/technical decision |
| REQ-008 | 5.7 | MRI when applicable | GAP-023 | Draft proposed / NOT approved | Not provided | Applicability decision pending; effective approved record and actual traceable test evidence as applicable | Open — evidence/technical decision |
| REQ-009 | 5.8 | Sterilization SAL 10^-6 | GAP-017 | Draft proposed / NOT approved | Not provided | WI references sterilization, validation evidence pending; effective approved record and actual traceable test evidence as applicable | Open — evidence/technical decision |
| REQ-010 | 6.1 | Kink stability | GAP-015 | Draft proposed / NOT approved | Not provided | Clinical limit risk-dependent; effective approved record and actual traceable test evidence as applicable | Open — evidence/technical decision |
| REQ-011 | 6.2 + Annex B | Corrosion metallic components | GAP-009 | Draft proposed / NOT approved | Not provided | Annex named; execution conditions/evidence missing; effective approved record and actual traceable test evidence as applicable | Open — evidence/technical decision |
| REQ-012 | 6.3 + Annex C | Negative-pressure deformation | GAP-008 | Draft proposed / NOT approved | Not provided | Annex named; max rated pressure/method missing; effective approved record and actual traceable test evidence as applicable | Open — evidence/technical decision |
| REQ-013 | 6.4.1 / Table 1 + Annex D | External connector tensile | GAP-010 | Draft proposed / NOT approved | Not provided | Quantified limits absent; effective approved record and actual traceable test evidence as applicable | Open — evidence/technical decision |
| REQ-014 | 6.4.2 / Table 2 + Annex E | Tubular/junction tensile | GAP-011 | Draft proposed / NOT approved | Not provided | Annex named; limits absent; effective approved record and actual traceable test evidence as applicable | Open — evidence/technical decision |
| REQ-015 | 6.5 + Annex F | Collection device impact | GAP-012 | Draft proposed / NOT approved | Not provided | No test link; effective approved record and actual traceable test evidence as applicable | Open — evidence/technical decision |
| REQ-016 | 6.6 + Annex G | Flow rate if claimed | GAP-014 | Draft proposed / NOT approved | Not provided | Claim applicability pending; effective approved record and actual traceable test evidence as applicable | Open — evidence/technical decision |
| REQ-017 | 6.7 | Retention means other than balloon | GAP-015 | Draft proposed / NOT approved | Not provided | Applicability based on design; no retention feature shown; effective approved record and actual traceable test evidence as applicable | Open — evidence/technical decision |
| REQ-018 | 6.8–6.10 + Annex I–L | Catheter balloon features | N/A pending design proof | N/A — provisional | Not provided | Controlled product-config/BOM proof of no catheter balloon; QA applicability authorization | Pending applicability proof |
| REQ-019 | 6.11 + Annex M | Aspiration/vacuum leak | GAP-013 / GAP-022 | Draft proposed / NOT approved | Not provided | Manual screen not enough to confirm method; effective approved record and actual traceable test evidence as applicable | Open — evidence/technical decision |
| REQ-020 | 7.1–7.2 | Manufacturer marking/packaging | GAP-016 / GAP-018 / GAP-022 | Draft proposed / NOT approved | Not provided | Partial WI; label proof missing; effective approved record and actual traceable test evidence as applicable | Open — evidence/technical decision |
| REQ-021 | 7.3 | Separate IFU (if supplied) | GAP-019 | Draft proposed / NOT approved | Not provided | IFU source unavailable; effective approved record and actual traceable test evidence as applicable | Open — evidence/technical decision |

# Detailed GAP closure and ownership

## GAP-001 | P2 — Unreconciled controlled document identity

ISO says / source: Document-control traceability issue (not a discrete ISO 20697 product characteristic). Effective controlled instructions and forms must be established before use.

Required objective evidence: Approved document master list; historical supersession/revision records; signed approved forms and WI; traceability reconciliation. Status: Open — Approval; no evidence supplied.

Control route: Document control | Responsible function: QA / Document Control.

Available executed evidence: None supplied for verification. Draft documentary action is described in companion revised WI and forms; not approved. Closure: Open — Approval.

## GAP-002 | P1 — Intended performance and risk file are not linked

ISO says / source: §4 ISO p.3 (PDF p.9): drainage system must maintain drainage; catheter must access target safely. §5.1 ISO pp.3–4 (PDF pp.9–10): apply design risk-management process and verify via risk-management file.

Required objective evidence: Approved risk analysis and intended-performance verification evidence; product configuration / variant mapping. Status: Open — Evidence.

Control route: Design qualification / risk management | Responsible function: QA / Engineering.

Available executed evidence: None supplied for verification. Draft documentary action is described in companion revised WI and forms; not approved. Closure: Open — Evidence / Approval.

## GAP-003 | P1 — Biocompatibility evidence linkage absent

ISO says / source: §5.2 ISO p.4 (PDF p.10): the device must be free from biological hazards based on appropriate ISO 10993-1 evaluation and testing.

Required objective evidence: Signed biological evaluation / test reports as justified, material/supplier configuration and change-control records. Status: Open — Evidence.

Control route: Biological evaluation / change control | Responsible function: QA / Regulatory / Engineering.

Available executed evidence: None supplied for verification. Draft documentary action is described in companion revised WI and forms; not approved. Closure: Open — Evidence / Approval.

## GAP-004 | P2 — Detectability status not governed by risk decision

ISO says / source: §5.3 ISO p.4 (PDF p.10): X-ray or alternative detectability of effective catheter length is required IF risk assessment determines so; ASTM F640 is an example method, not universally mandatory.

Required objective evidence: Approved risk decision + effective-length detectability report (if required), link to variant and formulation. Status: Open — Evidence.

Control route: Design qualification / routine appearance check | Responsible function: QA / Engineering.

Available executed evidence: None supplied for verification. Draft documentary action is described in companion revised WI and forms; not approved. Closure: Open — Evidence / Approval.

## GAP-005 | P2 — Surface finish method does not reflect risk-based magnification

ISO says / source: §5.4 ISO p.4 (PDF p.10): external effective catheter surface must be free of extraneous matter and defects posing unacceptable harm risk; minimum 2.5× magnification only if risk assessment deems it necessary.

Required objective evidence: Approved visual defect standard; risk-based magnification decision; completed line inspection records; reject traceability. Status: Open — Technical Input / Evidence.

Control route: Routine in-process inspection | Responsible function: QC / QA.

Available executed evidence: None supplied for verification. Draft documentary action is described in companion revised WI and forms; not approved. Closure: Open — Technical Input / Evidence.

## GAP-006 | P2 — Catheter size designation and OD tolerances not controlled

ISO says / source: §5.5.2 ISO p.4 (PDF p.10): nominal outer diameter designated in mm, rounded upwards to 0.1 mm; stated OD tolerance ±1 French; non-round by maximum axis. §5.5.3 p.4: effective length expressed in mm (<100 mm), or mm/cm (≥100 mm); ISO does not set length tolerance.

Required objective evidence: Approved drawings by size; dimensional control plan; calibrated gauge records; actual inspection results. Status: Open — Technical Input / Evidence.

Control route: Routine in-process dimensional checking | Responsible function: Engineering / QC.

Available executed evidence: None supplied for verification. Draft documentary action is described in companion revised WI and forms; not approved. Closure: Open — Technical Input / Evidence.

## GAP-007 | P1 — Misconnection hazards lack verified risk review

ISO says / source: §5.6 ISO p.5 (PDF p.11): avoid connector misconnection based on manufacturer risk assessment following ISO 80369-1 general principles; ISO 20697 does not prescribe one universal connector design.

Required objective evidence: Approved connector compatibility and misconnection risk analysis, design-verification evidence and completed SKU-level fit checks. Status: Open — Evidence.

Control route: Design risk qualification + routine component identity | Responsible function: Engineering / QA / QC.

Available executed evidence: None supplied for verification. Draft documentary action is described in companion revised WI and forms; not approved. Closure: Open — Evidence / Approval.

## GAP-008 | P1 — Maximum vacuum deformation test not operationally defined

ISO says / source: §6.3 ISO p.6 (PDF p.12), Annex C ISO p.14 (PDF p.20): device under manufacturer-defined maximum negative pressure must not collapse to impair function; ready-to-use test; body-inserted parts (37±2)°C ≥15 min, others (22±2)°C and RH ≥40% ≥15 min; apply max pressure ≥60 s.

Required objective evidence: Approved max negative pressure and completed Annex C report incl conditions, duration, finding and functional acceptance, mapped to SKU. Status: Open — Technical Input / Evidence.

Control route: Qualification / approved periodic plan | Responsible function: Engineering / Laboratory / QA.

Available executed evidence: None supplied for verification. Draft documentary action is described in companion revised WI and forms; not approved. Closure: Open — Technical Input / Evidence.

## GAP-009 | P2 — Corrosion Annex B procedure and evidence unspecified

ISO says / source: §6.2 ISO p.6 (PDF p.12) conditional on relevant exposed metal corrosion risk; Annex B ISO p.13 (PDF p.19): 0.15 mol/L NaCl 22±5°C 5 h; boiling water 30 min; 37±2°C 48 h; dry and inspect for corrosion.

Required objective evidence: Approved exposed-metal risk decision, materials trace, Annex B completed report and incoming rust checks; Status: Open — Evidence.

Control route: Incoming inspection + qualification / supplier verification | Responsible function: Supplier Quality / Laboratory / QC.

Available executed evidence: None supplied for verification. Draft documentary action is described in companion revised WI and forms; not approved. Closure: Open — Evidence / Approval.

## GAP-010 | P1 — External connector tensile limits absent

ISO says / source: §6.4.1 ISO p.6 (PDF p.12), Table 1 / Annex D ISO p.15 (PDF p.21): for outer diameter ≥2≤4 mm minimum 5 N; >4 mm minimum 15 N; <2 mm manufacturer determines via risk. Annex D tensile speed 500 mm/min; no separation.

Required objective evidence: Verified Annex D report with force, OD, test speed and separation outcome + controlled part/config mapping. Status: Open — Evidence.

Control route: Type qualification / periodic verification | Responsible function: Laboratory / Engineering / QA.

Available executed evidence: None supplied for verification. Draft documentary action is described in companion revised WI and forms; not approved. Closure: Open — Evidence / Approval.

## GAP-011 | P1 — Tube and bonded-junction tensile criteria missing

ISO says / source: §6.4.2 ISO p.7 (PDF p.13), Table 2 / Annex E ISO pp.16–17 (PDF pp.22–23): minimum 10 N for test piece OD 2–4 mm and 20 N for >4 mm; <2 mm risk-based; test each relevant tubular portion/junction with specified conditioning and strain rate.

Required objective evidence: Completed Annex E force reports for each relevant segment/junction, calibrated tester and approved sampling/control plan. Status: Open — Evidence / Technical Input.

Control route: Qualification + approved periodic verification | Responsible function: Engineering / Laboratory / QC.

Available executed evidence: None supplied for verification. Draft documentary action is described in companion revised WI and forms; not approved. Closure: Open — Technical Input / Evidence.

## GAP-012 | P1 — Bellow collection-device impact performance not substantiated

ISO says / source: §6.5 ISO p.7 (PDF p.13), Annex F ISO pp.18–19 (PDF pp.24–25): collection device must not leak and suction source must not lose >2% vacuum in defined impact test; drop height 700 mm at (22±5)°C, measure pressure before and ≥60 s after drop.

Required objective evidence: Annex F report, equipment qualification, before-after readings, SKU match and QA approval. Status: Open — Evidence.

Control route: Type / qualification testing | Responsible function: Engineering / Laboratory / QA.

Available executed evidence: None supplied for verification. Draft documentary action is described in companion revised WI and forms; not approved. Closure: Open — Evidence / Approval.

## GAP-013 | P1 — Vacuum air-leak testing lacks ISO method and limits

ISO says / source: §6.11 ISO p.8 (PDF p.14), Annex M ISO pp.35–36 (PDF pp.41–42): no air ingress at manufacturer-defined max negative pressure; de-aerated water 22±5°C; stabilize 120 s then maintain 120 s; record location of any leakage.

Required objective evidence: Approved vacuum-pressure specification; ISO Annex M test report showing 120+120 s, pressure and outcome; controlled lot/variant mapping. Status: Open — Technical Input / Evidence.

Control route: Routine manual check + qualification/periodic Annex M | Responsible function: Laboratory / Engineering / QC.

Available executed evidence: None supplied for verification. Draft documentary action is described in companion revised WI and forms; not approved. Closure: Open — Technical Input / Evidence.

## GAP-014 | P2 — Claimed flow-rate verification is conditional but unmapped

ISO says / source: §6.6 ISO p.7 (PDF p.13), Annex G ISO pp.20–21 (PDF pp.26–27): IF flow rate claimed, each lumen ≥80% of stated for nominal OD <1.0 mm, or ≥90% for OD ≥1.0 mm; three measurements per lumen, average ml/min under Annex G hydrostatic conditions.

Required objective evidence: Approved claim assessment plus Annex G results if claimed; current product config mapping. Status: Open — Applicability / Evidence.

Control route: Conditional design qualification | Responsible function: Engineering / Regulatory / Laboratory.

Available executed evidence: None supplied for verification. Draft documentary action is described in companion revised WI and forms; not approved. Closure: Open — Technical Input / Evidence.

## GAP-015 | P2 — Kink and retention requirements lack applicability decisions

ISO says / source: §6.1 ISO p.6 (PDF p.12): manufacturer determines clinically relevant kink value from intended use and risk; ISO does not mandate numeric minimum. §6.7 p.7 (PDF p.13): if non-balloon retention means exist, demonstrate prevention of undesired displacement; Annexes A/H are informative.

Required objective evidence: Approved risk-based clinical kink/retention criteria and verification, or signed documented N/A. Status: Open — Technical Input.

Control route: Risk-based design qualification | Responsible function: Engineering / QA.

Available executed evidence: None supplied for verification. Draft documentary action is described in companion revised WI and forms; not approved. Closure: Open — Technical Input.

## GAP-016 | P2 — Bellows collection capacity controlled as options, not measured limits

ISO says / source: §7.2(d) ISO p.9 (PDF p.15): effective collection capacity expressed in ml on packaging where collection device present. §4 p.3: device must maintain intended drainage. ISO does not impose one universal capacity tolerance.

Required objective evidence: Approved capacity tolerances, validated practical method, calibrated cylinder and completed line measurements. Status: Open — Technical Input / Evidence.

Control route: Routine in-process approved sampling | Responsible function: Engineering / QC.

Available executed evidence: None supplied for verification. Draft documentary action is described in companion revised WI and forms; not approved. Closure: Open — Technical Input / Evidence.

## GAP-017 | P1 — Sterile-product validation release traceability incomplete

ISO says / source: §5.8 ISO p.6 (PDF p.12): sterile drainage catheter/accessories comply with applicable sterilization standards and demonstrate SAL 10^-6.

Required objective evidence: Approved SAL validation/revalidation, sterilization batch record, sterile release and lab evidence. Status: Open — Evidence; no executed records supplied.

Control route: Validated process + each sterilization load release | Responsible function: QA / Sterilization / QC.

Available executed evidence: None supplied for verification. Draft documentary action is described in companion revised WI and forms; not approved. Closure: Open — Evidence / Approval.

## GAP-018 | P2 — Device/package marking details not fully traceable

ISO says / source: §7.1 and §7.2 ISO p.9 (PDF p.15): marking legible, durable; packaging minimums include special storage/handling, single-use indicator, detectability where appropriate, collection capacity, pre-evacuated vacuum stability, rated max negative pressure and balloon volume where applicable.

Required objective evidence: Approved label artwork, marking durability record, approved storage/pressure claims, completed packaging form. Status: Open — Evidence / Technical Input.

Control route: Packaging inspection + design evidence | Responsible function: Regulatory / QA / Packaging.

Available executed evidence: None supplied for verification. Draft documentary action is described in companion revised WI and forms; not approved. Closure: Open — Technical Input / Evidence.

## GAP-019 | P2 — IFU mandatory content verification not referenced

ISO says / source: §7.3 ISO pp.9–10 (PDF pp.15–16): if separate IFU supplied, describe intended use when not evident, correct device combinations, warnings/precautions, issue/revision, applicable single-use, MRI, additive/coating and retention instructions.

Required objective evidence: Approved IFU, clause-by-clause label/IFU checklist and change-control record. Status: Open — Evidence.

Control route: Document design verification / artwork review | Responsible function: Regulatory / QA.

Available executed evidence: None supplied for verification. Draft documentary action is described in companion revised WI and forms; not approved. Closure: Open — Evidence / Approval.

## GAP-020 | P2 — In-process results are not uniquely quantified or traceable

ISO says / source: ISO §4 and §§5.4, 5.5, 6.4, 6.11 (pp.3–8; PDF pp.9–14) require appropriate verified functional/performance controls. The need for completed, traceable records is a company quality-control/documentation issue; ISO 20697 does not itself prescribe this form layout.

Required objective evidence: Approved control plan and AQL/sampling, executed forms with measured values, calibrated instruments and handling of fails. Status: Open — Technical Input / Evidence.

Control route: Routine in-process control | Responsible function: QC / Production / QA.

Available executed evidence: None supplied for verification. Draft documentary action is described in companion revised WI and forms; not approved. Closure: Open — Technical Input / Evidence.

## GAP-021 | P2 — Multi-channel table dimensions and revision are ambiguous

ISO says / source: §5.5.2–5.5.3 ISO p.4 (PDF p.10) provide OD designation principles; ISO does not define the product-specific silicone channel table dimensions. This is a company drawing and controlled-form clarity issue.

Required objective evidence: Approved drawing identifying all OD columns and size tolerances; approved corrected controlled form; completed records and revision reconciliation. Status: Open — Technical Input / Approval.

Control route: Routine extrusion in-process checks | Responsible function: Engineering / QC / Document Control.

Available executed evidence: None supplied for verification. Draft documentary action is described in companion revised WI and forms; not approved. Closure: Open — Technical Input / Approval.

## GAP-022 | P2 — Vacuum stability and rated pressure claim mapping missing

ISO says / source: §7.2(e)–(f) ISO p.9 (PDF p.15): pre-evacuated suction source requires a stability date tied to retention of ≥80% initial negative pressure and maximum pressure rating in Pa where appropriate; §6.3/6.11 use manufacturer maximum pressure for tests.

Required objective evidence: Approved pressure spec, pre-evacuation evidence, storage/stability report, label and test linkage. Status: Open — Technical Input / Evidence.

Control route: Design qualification / shelf-life validation | Responsible function: Engineering / QA / Regulatory.

Available executed evidence: None supplied for verification. Draft documentary action is described in companion revised WI and forms; not approved. Closure: Open — Technical Input / Evidence.

## GAP-023 | P2 — MRI-risk applicability has not been documented

ISO says / source: §5.7 ISO p.6 (PDF p.12): if applicable, evaluate hazards for device and accessories in MRI environment by appropriate method.

Required objective evidence: Approved MRI risk determination and evidence/label review when applicable. Status: Open — Applicability.

Control route: Risk-based design review | Responsible function: QA / Regulatory / Engineering.

Available executed evidence: None supplied for verification. Draft documentary action is described in companion revised WI and forms; not approved. Closure: Open — Technical Input.

# Technical Inputs Required Before Operational Implementation

Approved true WI/form revisions and supersession map.

Approved SKU/BOM/drawings, tube OD-ID/channel definition, dimensions, company effective length and bellows capacity tolerances.

Maximum stated negative pressure, pressure claims and vacuum stability date/validation for pre-evacuated variants.

ISO applicability decisions: flow claims, catheter detectability, MRI, kink value, retention means and whether any actual catheter balloon exists.

Approved AQL / lot sampling / shift frequencies, proper in-process leak method, glue/cure requirements and gauge methods.

Executed and approved required ISO Annex test reports, risk file, biological evaluation, sterilization SAL and packaging/IFU content.

# Manager / QA Authorization

QC Manager decision:  [ ] Request technical information  [ ] Approve revision for formal document-control review  [ ] Return for rework

QC Manager name/signature/date: _____________________   QA approval: _____________________   Document Control effective revision/date: _____________________

Record of evidence review date and auditor: ___________________________     Next approved closure review: ___________________________