QUALITY SYSTEMS / CONTROLLED DRAFT

# Revised IN-PROCESS Control Forms

F-823-i12 and F-823-i108 • Not a final Inspection & Test Report

Wound Drainage Set   |   ISO 20697:2018(E), first edition 2018-06, corrected 2018-09

Status: DRAFT FOR QC/QA/ENGINEERING REVIEW   •   No actual product conformity or ISO closure is asserted.

Source documents: WI-8-2-2-23; F-823-i12; F-823-i108; ISO 20697:2018(E). Evidence supplied consists of unexecuted document templates, not completed production/test records.

# Purpose and use

This file fulfills the requested editable “Revised Inspection Report” deliverable as correctly classified IN-PROCESS CONTROL FORMS. Original F-823-i12 and F-823-i108 rows are retained in redesigned checklists. Actual pre-sterilization process observations are recorded here; ISO Annex qualification and final release records are linked, not duplicated as universal hourly lot tests. All proposed form revisions require controlled approval before use.

# Form A — Wound Drainage Set | F-823-i12 (Revised Draft)

| Header field | Entry |
| --- | --- |
| Form revision (TBA) | ____________________________________ |
| WI revision (TBA) | ____________________________________ |
| SKU / production order | ____________________________________ |
| Product size / nominal bellows capacity | ____________________________________ |
| Lot / batch / line / shift | ____________________________________ |
| Date / inspection time / sample ID | ____________________________________ |
| Machine/station / operator / QC | ____________________________________ |
| Approved control plan / drawing reference | ____________________________________ |
| Instrument(s), gauge IDs / calibration status | ____________________________________ |

| No. | Original in-process characteristic (retained) | Revised company/ISO-aligned criterion | Method |
| --- | --- | --- | --- |
| 01 | Tube / bellow pump clarity | Clear, approved appearance | Visual |
| 02 | Tube gluing / insertion | Correct assembly; no manual pullout | Visual/manual + evidence link |
| 03 | Tube with slide clamp | Present when specified by SKU | Visual |
| 04 | Fittings status | Correct assembly, no damage | Visual |
| 05 | Leak screen for bellows | No observed leakage by approved in-process method | Visual/manual |
| 06 | Bellows capacity | 200 / 400 / 600 ml as per SKU; tolerance OPEN | Calibrated volume measurement |
| 07 | Connector type | Matches approved BOM and intended connection | Visual/drawing |
| 08 | Trocar type / size | Matches approved SKU/drawing | Visual/caliper if required |
| 09 | Trocar sharpness | Sharp, protected, undamaged | Visual / approved safe method |
| 10 | Redon drainage tube effective length | 500 mm with multiple holes only where drawing requires; tolerance OPEN | Measurement |
| 11 | Bellows graduation scale | Clear legible graduations per artwork | Visual |
| 12 | Punching holes quality | No fiber, obstruction or roughness | Visual |
| 13 | Connector–bellows attachment | Correct fit; no manual pullout (see Annex D/E proof) | Visual/manual |
| 14 | Correct component specification | Matches approved released drawing | Document/visual |
| 15 | Drilled hole on bellows neck | Unobstructed, no fiber | Visual |
| 16 | Apply rubber on hole | Yes/No as defined for SKU | Visual/BOM |
| 17 | Hanger pieces | Present where specified | Visual |

# Form A — Reusable Actual Result Log (record multiple rows per shift)

Each row records one performed control, including actual numerical measurement when required. Duplicate blank rows electronically or in Word as needed; original hourly grid 7,8,9,10,11,12,1,2,3,4,5,6 may be used as the scheduling reference after frequency approval.

| Date/time | Item No. / Sample ID | Actual result and unit / observation | Approved limit or spec rev | Instrument ID | P/F/N/A | NCR / action |
| --- | --- | --- | --- | --- | --- | --- |
| _____ | _____ | ________________ | ________ | ______ | _____ | _____ |
| _____ | _____ | ________________ | ________ | ______ | _____ | _____ |
| _____ | _____ | ________________ | ________ | ______ | _____ | _____ |
| _____ | _____ | ________________ | ________ | ______ | _____ | _____ |
| _____ | _____ | ________________ | ________ | ______ | _____ | _____ |
| _____ | _____ | ________________ | ________ | ______ | _____ | _____ |
| _____ | _____ | ________________ | ________ | ______ | _____ | _____ |
| _____ | _____ | ________________ | ________ | ______ | _____ | _____ |
| _____ | _____ | ________________ | ________ | ______ | _____ | _____ |
| _____ | _____ | ________________ | ________ | ______ | _____ | _____ |
| _____ | _____ | ________________ | ________ | ______ | _____ | _____ |

# Form B — Silicone Multi Channel Tube | F-823-i108 (Revised Draft)

| Header field | Entry |
| --- | --- |
| Form revision: resolve filename Rev.6 versus sheet Rev.4 | ____________________________________ |
| Product SKU / CH / tube arrangement | ____________________________________ |
| Trocar size / drawing revision | ____________________________________ |
| Lot / order / date / shift | ____________________________________ |
| Extrusion line / operator | ____________________________________ |
| QC inspector / QC reviewer | ____________________________________ |
| Control plan / approved frequency | ____________________________________ |
| Instrument IDs / calibration status | ____________________________________ |

| No. | Original silicone in-process characteristic | Revised requirement | Method |
| --- | --- | --- | --- |
| 01 | Clarity / appearance | Smooth, flexible and free of identified critical defects | Visual |
| 02 | Tube quality status | Unobstructed; no unacceptable discoloration | Visual |
| 03 | Total length of tube | Per approved component drawing; record actual mm | Measurement |
| 04 | Both tubing tips after cut | Smooth, correct cut both ends | Visual |
| 05 | No irregularities | No bumps/black spots/fish eyes/sink marks | Visual |
| 06 | Tube dimension | OD, ID, segment lengths per drawing with tolerances | Caliper/approved gauge |
| 07 | Multi-channel quality | Required channels (source: four), correct openings/cutting | Visual/measurement |
| 08 | Barium sulfate status | Approved blue/white x-ray line workmanship; separate detectability evidence | Visual |
| 09 | Trocar fit on channel tubing | Correct attachment/no manual pullout; separate tensile proof | Visual/manual |
| 10 | Trocar dimension | Actual dimensions per drawing / approved limits | Caliper/approved gauge |

# Form B — Size Reference (source values only, unapproved until drawing reconciliation)

The following numbers were copied as design-reference data from source F-823-i108 rows 21–26. Historical columns labelled “OD mm” twice: their engineering meaning is OPEN TECHNICAL INPUT REQUIRED. No tolerance was inferred.

| CH | OD×ID source | OD column E | OD column G | A mm | B mm | C mm | Total mm |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 10 | 3.3×1.3 | 3.8 | 3.3 | 800 | 100 | 300 | 1200 |
| 15 | 5.0×3.0 | 5.5 | 5.0 | 800 | 100 | 300 | 1200 |
| 18 | 5.5×3.5 | 6.0 | 6.0 | 800 | 100 | 300 | 1200 |
| 19 | 6.0×4.0 | 6.5 | 6.5 | 800 | 100 | 300 | 1200 |
| 24 | 8.0×5.2 | 8.5 | 8.0 | 800 | 100 | 300 | 1200 |

# Form B — Numeric/Visual Result Log

Reference time positions in original template are 7,9,11,1,3,5 (two-hour grid). Approved actual control plan governs inspection frequency.

| Time | Item/sample ID | Location / OD-ID channel | Acceptance / drawing | Actual value & unit | Instrument | P/F/N/A |
| --- | --- | --- | --- | --- | --- | --- |
| _____ | _____ | ________________ | ____________ | ____________ | ______ | _____ |
| _____ | _____ | ________________ | ____________ | ____________ | ______ | _____ |
| _____ | _____ | ________________ | ____________ | ____________ | ______ | _____ |
| _____ | _____ | ________________ | ____________ | ____________ | ______ | _____ |
| _____ | _____ | ________________ | ____________ | ____________ | ______ | _____ |
| _____ | _____ | ________________ | ____________ | ____________ | ______ | _____ |
| _____ | _____ | ________________ | ____________ | ____________ | ______ | _____ |
| _____ | _____ | ________________ | ____________ | ____________ | ______ | _____ |
| _____ | _____ | ________________ | ____________ | ____________ | ______ | _____ |
| _____ | _____ | ________________ | ____________ | ____________ | ______ | _____ |
| _____ | _____ | ________________ | ____________ | ____________ | ______ | _____ |

# Form C — Non-Routine Qualification and External Evidence Link

Complete only where relevant/required to document current approved evidence. No routine Annex testing is automatically scheduled by this table.

| Requirement | Evidence method / subject | Function | Approved report ID | Status |
| --- | --- | --- | --- | --- |
| ISO 5.1 / 5.2 | Risk management & biological evaluation | QA / Engineering | Pending | Open |
| ISO 5.3 / 5.6 / 5.7 | Detectability, connector risk, MRI if applicable | Engineering / QA | Pending | Open |
| 6.2 / Annex B | Trocar corrosion, risk conditional | Lab / supplier quality | Pending | Open |
| 6.3 / Annex C | Pressure deformation with manufacturer max pressure | Lab / Engineering | Pending | Open |
| 6.4.1 / Annex D | External connector tensile (Table 1) | Lab / Engineering | Pending | Open |
| 6.4.2 / Annex E | Catheter/junction tensile (Table 2) | Lab / Engineering | Pending | Open |
| 6.5 / Annex F | Bellows impact / vacuum loss | Lab / Engineering | Pending | Open |
| 6.6 / Annex G | Claimed flow rate (conditional) | Lab / Regulatory | Pending | Open |
| 6.11 / Annex M | Aspiration air ingress/vacuum leak | Lab / Engineering | Pending | Open |
| 5.8 / 7.1–7.3 | Sterilization, label/IFU, final authorized release | QA / Sterilization | Pending | Open |

# Sign-Off and Disposition (for in-process production stage only)

| Role | Name / signature / date | Decision |
| --- | --- | --- |
| Operator / production | ________________________ | Data recorded; no authorization to release finished goods |
| QC Inspector | ________________________ | PASS / HOLD / NCR: __________ |
| QC Reviewer / Supervisor | ________________________ | Reviewed: __________ |
| Document Control (form approval only) | ________________________ | Form effective revision: ______ |

Final sterile release decision is outside this form. Refer to WI-7-5-1-1, packaging form F-823-i4-2 and F-8-2-4 as applicable. Current forms are unapproved drafts.